[image: image2.png]R e e N Huazheng




[image: image7.png]Huazheng





JZF-6 

Calibration pulse generator

 [image: image1.png]



Huazheng Electric Manufacturing (Baoding) Co., Ltd

Dear user:

Thank you for choosing JZF-6Calibration pulse generator.

     We hope that this instrument can make your work easier and more enjoyable, so that you can get the feeling of office automation in the test and analysis work.

  Before using the instrument, please read this manual, and operate and maintain the instrument according to the manual to prolong its service life.

"Just a light press, the test will be completed automatically" is the operating characteristics of this instrument.

     If you are satisfied with this instrument, please tell your colleagues; if you are not satisfied with this instrument, please call (0312) 6775656 to tell you to serve you at all times-Baoding Huazheng Electric Manufacturing Co., Ltd., our company will definitely make you satisfied !

Contents
1I. Overview


1II. Specifications


1III. Instructions


2IV. Note


2V. Packing List




I. Overview
JZF-6 Calibration pulse generator is a small battery-powered partial discharge corrector. It has the characteristics of small size, easy to carry, convenient synchronization, wide calibration power output range, etc. By selecting different combinations of pulse voltage amplitude and injection capacitance, it can output multiple calibration values from 10pC to 5000pC, which can meet various types of test samples. Correction requirements. It is suitable for any kind of test circuit recommended by the International Electrotechnical Commission IEC-270.

II. Specifications

1. Output pulse amplitude file (U): three files 1, 2, 5;

2. Injection capacitor file (C): X10pF, X100pF, X1000pF;

3. The total output partial discharge correction amount (Q=U×C): 10pC~5000pC, error ±5%;

4. Pulse front: <60nS;

5. The back edge of the pulse:> 100uS;

6. Output pulse polarity: alternating positive and negative;

7. Repetition frequency: 1.2kHz;

8. Frequency variation range: >±100Hz;

9. Size: 160×120×50mm3;

10. Weight: 0.5kg;

11. Battery: 6F22 9V.

III. Instructions

1. Open the back cover of the JZF-6 calibration pulse generator and insert the battery before use. After closing the battery box cover, connect the red and black output terminals to the wires. The wire on the red terminal should be as short as possible and connected to the high-voltage end of the test product, and the wire of the black terminal should be connected to the low-voltage end of the test product.

2. Set the calibration power switch to any suitable gear to calibrate. The correction pC amount of the JZF-6 calibration pulse generator is determined by the product of the pulse amplitude and the output injection capacitance, namely: Q(pC) = U(V)×C(pF).

For example, if a calibration amount of 5000pC is required, turn the knob to the "5" position and select “1000pF” for the injection capacitor, and the output calibration amount is 5000pC.

The frequency can be adjusted around 1.2 kHz (correction pulse synchronization adjustment). The voltmeter on the panel indicates the battery voltage. Generally, it can be guaranteed to work if it is above 7V. Please pay attention to replace the battery when it is below 7V.

IV. Note

After calibration, be sure to disconnect the connection line between the JZF-6 calibration pulse generator and the test product. To prevent the JZF-6 calibration pulse generator from being damaged by high voltage.

V. Packing List
	No.
	Item
	Qty

	1
	Dedicated measuring line （red）
	1

	2
	Dedicated measuring line  （black)
	1



[image: image3.png]R e e N Huazheng



[image: image4.png]R e e N Huazheng



[image: image5.png]R e e N Huazheng



[image: image6.png]R e e N Huazheng



[image: image7.png]