[image: image37.png]R e e N Huazheng




[image: image45.png]U~ U~
[ [ e b
add voltage add voltage
on ab phase on ca phase





HZ-6300C  

Transformer Capacity Testing Instrument

[image: image1.png]



Huazheng Electric Manufacturing(Baoding) Co.,Ltd
Dear user:

Thank you for choosing HZ-6300C Transformer Capacity Testing Instrument.

     We hope that this instrument can make your work easier and more enjoyable, so that you can get the feeling of office automation in the test and analysis work.

  Before using the instrument, please read this manual, and operate and maintain the instrument according to the manual to prolong its service life.

"Just a light press, the test will be completed automatically" is the operating characteristics of this instrument.

     If you are satisfied with this instrument, please tell your colleagues; if you are not satisfied with this instrument, please call (0312) 6775656 to tell you to serve you at all times-Baoding Huazheng Electric Manufacturing Co., Ltd., our company will definitely make you satisfied !

Content

1I.Introduction


2II.Definition of Testing Transformer


4III.Technical Parameters


4IV.Function of Interface


19V.wiring method


29VI. Note


29VII.Packing List


30Appendix-Catalogue


33B  GB/T 6451 and GB/T 10228 Appendix B GB/T 6451 and GB/T 10228





I.Introduction
Transformer Capacity Testing Instrument is specialized in bad power users evading basic electricity secretly adds capacity problems and is a high precision instrument specially used for the measurement of the transformer capacity and parameters. This product uses color LCD monitor, English display, simple and quick menu operation, man-machine interface is friendly. The internal of the instrument has a rechargeable battery and it can work without external power source. A fully charged battery can be continuously measured more than 100 sets. In addition, the instrument can provide three-phase precise 50Hz sine wave AC test power without the external three-phase test power, the pressure regulator, up flow reactor and other auxiliary equipment in measuring the capacity and transformer load loss, thus greatly improving the efficiency of the work.

This is a multi-functional measuring instrument, which is equivalent to two common devices: activity transformer capacity tester and transformer parameter tester. It can be used for test transformer capacity, model, no-load current, no-load loss, load loss, impendence voltage and so on.

This product has many advantages such as small, light , high accuracy, stability, simple operation and so on. It can replace the multi-meter method to measure the transformer loss and capacity with simple wiring and convenient testing so that the work efficiency can be greatly improved.

Function

We can measure a variety of distribution transformer, power transformer capacity instead of the additional power. It is more convenient and efficient.

Combined with the power, voltage, boost up flow and other equipment, it measures various transformer no-load current, no-load loss, impedance voltage, load loss and so on.

All results are automatically carried out relevant correction. You only need to input the related parameters (such as temperature, load adjustment coefficients), then the instrument can automatically correct such as waveform deformity correction, temperature correction, correction of non rated voltage, rated current correction.

The maximum voltage is up to 750V and the maximum current is up to 100A. It is equipped with protection circuit. 

The voltage and current range can be very flexible and easy to expand.

It has large screen, high brightness LCD display, English menu and operation tips, friendly man-machine interface and wide temperature, brightness adjustable type LCD screen, which can adapt to the different temperature.

According to the test results of no-load and load loss, it can automatic presume the three-phase oil immersed distribution transformer and power transformer’s performance levels as the reference (this instrument is currently only available for three-phase oil immersed transformer and power transformer performance level code. Others are invalid.)

II.Definition of Testing Transformer 

Transformer capacity and characteristic test is the professional test. There are many setting part and items of testing method. 

Tap

    In the transformer with topped winding, every tap in this winding stands for a fixed turn number of windings in this tap, which also means a fixed turn number of windings between this tap winding with other constant windings 

Principal Tapping

Tap corresponds to the rated parameters. Attention: The data of load/no-load loss and no-load current is the data of transformer under the principal tapping (except for appointed other taps)

 On-load Tap Charger

   It is the device that is used to change the position of tap of transformer winding under operating with excitation or no-load. 

Off-circuit Tap Charger

    It is the device that is used to change the position of tap of transformer winding under operating without excitation.

No-load Loss (commonly known as: iron loss, iron loss)

   When rated voltage (tapped voltage) under rated frequency is applied to terminal of one winding and open circuit of other windings, it is the active power.

No-load Current

When rated voltage (tapped voltage) under rated frequency is applied to terminal of one winding and open circuit of other windings, it is the root average of current.

Note 1: For three phase transformer, it is the average of current of three phase terminals.

Note 2: Generally use the percent of current to display. For the multi-winding transformer, it is based on the winding with max rated capacity.

Load Loss (Short Circuit Loss/Copper Loss)

It is the active power in one pair of winding, when rated current flows through the line terminal and short circuit of the other terminal under the rated frequency and reference temperature. If it has other windings, please make them open circuit.

Note1: For the double winding of transformer, there is just one pair of windings and one value of load loss.

       For multi-windings of transformer, there are multi pairs and value of loading loss.

Note2: When the capacity of the two windings, it should be based on the rated current of the windings with the lower rated capacity, and appointed to the reference capacity.

Total Loss

    The sum of load loss and no-load loss.

Short Circuit Impedance

Generally, short circuit impedance is the impedance of the pair winding.

Under the rated frequency and reference temperature, the equivalent series impedance of one pair windings between the terminals of winding is 

Z=R+jX(Ω). Now, the terminal of the other one need to short circuit. If there are other windings, they need to open circuit.

Note1: For the transformer with tapped windings, short circuit impedance need to appoint the tap, if there is no special instruction, it is the principal tap. For the three phase transformer, it means the impedance of single phase.

Note 2: It can be represented using non-dimensional relative value that means fraction Z of reference impedance Zref in same winding of one pair windings, as follows:[image: image38.png]R e e N Huazheng




[image: image39.png]R e e N Huazheng




Formula is suitable for the three phase and single phase transformer.

U——Voltage of Z and Zref in windings（rated voltage or tapped voltage）

Sr——Basic value of rated capacity.

The relative value is equal to the ratio between the voltage and rated voltage (tapped voltage) when operator wants to get the rated current (tapped current) in short circuit test. This voltage also called the short-circuit voltage of this pair windings.  It is generally expressed as a percentage. 

III.Technical Parameters

1. The output range of built-in battery

   Voltage：0～10V

   Current：0～10A

2. The output range of characteristic test battery
   Voltage：0～750V

   Current：0～100A

3. Accurate
   Voltage, current：±0.2%

   Power：±0.5%（CosΦ>0.1）, ±1.0%（0.02<CosΦ<0.1）

Transformer capacity test range

6.3~125,000 KVA.

Working temperature
   -20℃～+60℃

6. The requirement of charger
    Electric supply AC160V～265V

7. Insulation
    ⑴ Insulation resistance between capacity , voltage, current and cover ≥100MΩ

    ⑵ Power frequency is 2KV(effective value) between the input power and cover, testing time is 1 minute. 

8. Volume 

    32cm×25cm×12cm

9.  Weight

    2.5KGS                    
IV.Function of Interface

1. Operation panel diagram.

The operation panel like figure 1 : including printer , LCD, keyboard, the wiring hole of capacity test , binding post of characteristic test and communication interface.

[image: image2.png]HOEE
B{B]UE[GE]
HOE®
HOEH®
HOE®

ww





Figure 1 Display and operating Panel diagram

2. Main Menu

After open the power switch, automatic enter into the main menu (Figure 2) Main Menu includes: "Transf. Power", "No - load loss", "Load loss", "Data Browser", "System Settings" and other five. You can use the "↑", "↓", key to select the desired function and click "OK" to enter.

At the same time, the instrument according to the system's internal clock chip, accurately display the current time.

[image: image3.png]Data Browser

Transf. Power No-Load Loss Load Loss

System Setting





3. Transf. Power

3.1 Transformer Power Testing Setup

In the main menu, select the "Transf.Power" item and click "OK", then it enter into parameter setting interface before the test, total are 10 items(Figure 3).

Primary U: Before making the determination of transformer Power, you need to input the working voltage of the transformer correctly, which is the transformer primary rated voltage. Enter directly through keyboard, Unit: KV.

Secondary U: Before making the determination of transformer Power, you need to input the working voltage of the transformer correctly, which is the transformer secondary rated voltage. Enter directly through keyboard, Unit: KV.

The input data of the primary voltage and secondary voltage are not higher than 500kv, and if the input data is not included in those following voltage levels, the instrument will change "Transformer Type" into "non-standard transformers." When testing "non-standard transformer," you need enter the measured transformer "impedance voltage" to measuring transformer capacity accurately.

The following 12 transformer voltage levels is included by our instrument: 10kV / 0.4kV, 10kV / 3kV-6.3kV, 35kV / 0.4kV, 35kV / 3.5kV-11kV, 6kV / 0.4kV, 6kV / 3kV-3.15kV, 6.3kV / 0.4kV , 12 6.3kV / 3kV-3.15kV, 10.5kV / 0.4kV, 10.5kV / 3kV-6.3kV, 11kV / 0.4kV, 11kV / 3kV-6.3kV etc.( In front of “/ ” is the transformer primary rated voltage, after “ / “ is the secondary transformer rated voltage).

Transf. Types: Setting the type of transformer under test. Main parameters settings include "SJ (73) Transf.", "S7.S9 (11) Transf. ", "S7.S9 (11) Power", "S13 Transf.", "Coating CRDT", "unCoating CRDT", "NonStand. Transf."etc,7 alternatives. "Nonstand.Transf." includes non-distribution transformer special transformer and those transformers which their rated voltages are not within the voltage levels listed in above table. When selecting "S7.S9 (11) Transf." or "S7.S9 (11) Power", and when the capacity test is completed, the system will estimate the type of the measured transformer, depending on the load loss data measured by instrument. The determination result can be used for operator reference. 

Nomial Uz(%): If you want to the accurate result of capacity When testing the "NonStand. Transf." , you should input impedance voltage accurately. You can enter data directly using the number keys. When testing the other transformers except "NonStand. Transf." type. The instrument will display impedance voltage according to the rated voltage and transformer type. Under normal circumstances, operators can carry out test without modify data except when the test impedance voltage has large difference with impedance voltage marked on the transformer nameplate, it is recommended to change its impedance voltage data, to make it closer to the marked "Nomial Uz(%)" data., better for transformer capacity testing.

Testing Temp.: Capacity test required temperature correction, so we need to input current temperature. Generally inputting data is the data of under test transformer’s  back side temperature add 10 ℃. You can enter the data directly via the numeric keys.

[image: image4.png]Transformer Power Testing Setup

Primary U  10.0KV Tapping Step
Secondary U 0.4KV Nomial Power
Transf. Types  $7.59(11)Transf. Transf. No.

Nomial Uz(%) 409, Transformer user

Testing Temp.  20°C Testers

Two of 3Taps
100.0KVA
000001

David

Lucy

Lo/





Figure 3 Transformer Power Testing Setup

Tapping Step: the tap-changer position of the transformer under test. Distribution transformers usually have three tap stalls; two of them are standard taps. When capacity measuring, please keep the tap-changer position of the transformer under test same as the data set in this item. if under test transformer’s tap position instead of three, please set this item to 2 tap position, at the same time change the transformer’s tap-changer position to standard tap position before the capacity test. The data can be changed by clicking "←", "→" key.

Nominal Power: As a reference of the measurement result, here input the nominal capacity of the transformer measured, in order to compare with the measured capacity. This item can be input directly via the numeric keys.
Transf. No.: 6 digits of number in order to check and save, inputting through keyboard

The above items are set finished, and after the wiring is correct (see later for details), click "OK" button can perform capacity test.

3.2 Transformer Power Test Results

 Test results screen shown in Figure 4. It includes the following:

[image: image5.png]Transformer Power Test Results

Load Loss: 4.96KW Tested Uz%: 3.94%
CNS Loadloss: 5.10KW Nominal Uz%: 4.00%
LoadLoss(75°C) :  6.03KW Tested Type: S9(11) Transf.
LoadLoss Error: 18.24% TestedR: 7.2Q

Final T.Rower: 500KVA

| OK |Retry | Save | | Print | Exit| Back |





Figure 4 Transformer Power Test Results

Load Loss: under the current test conditions the actual measured load loss;

CNS LoadLoss: if archive measured capacity result, it will display the transformer load loss CNS(Chinese National Standard) data of archive position;

LoadLoss(75℃): the measured load loss correction to the rated load loss data obtained by rated experimental conditions;

LoadLoss Error :correction percentage of error between loss and GB loss;

Tested Uz(%): impedance voltage of measured transformer in the current test conditions;

Nominal Uz(%):  impedance voltage data used during the capacity test;

Final T.Power: when measured transformer capacity can be archived, the item displays archived capacity data. When measured transformer capacity cannot be archived, the item is not displayed;

Tested Power: the item displays the actual measurement capacity of the transformer. When measured capacity can be archived, the item is not displayed;

Tested Type: When set the parameter before capacity test, transformer type is set to "SJ.S7.S9 (11) distribution transform " or "S7.S9 (11) power converter", the system will display the estimated type of transformer under test, such as SJ distribution transformer, or S7 distribution transformer (or power converter) or S9 (11) distribution transformer (or power converter). For tester reference. When the measured capacity can not be archived, will show "----."

Tested R: Displaying the short circuit resistance of testing transformer understand calibrating up to 75℃ rated condition.

After testing, Click “Save”, testing result can be saved in instrument in order to check in future; click “Print”, testing result can be printed; click “End”, Screen returns to the part of setting up parameter.

4. No-Load Loss.

Please choose “No-Load Loss“on the menu, and click “ok”, and then setup the no-load test. No-load test is the test of AC power (including boosting voltage, regulating voltage, boosting current)

4.1 Setting 

Before testing, necessary parameters need to set up. Operator can set up the parameter by moving “↑” and“↓” on the screen.

Related parameter as follows: primary voltage, secondary voltage, nominal capacity, transformer model, transformer connection, testing model, voltage ratio, current ratio, and transformer No., operator and so on. ” Transformer model, connection of transformer, testing model can set up by moving “↑” and“↓” , other can input data.

Primary U: Inputting the working voltage before testing is the primary rated voltage, which can input directing by using keyboard. Unit is KV.

[image: image6.png]Transformer No-LoadLoss Testing Setup

Primary U 10.00KV Testing Model 3P4W No-Load

Secondary U 0.40KV PTRatio 10

Nomial Power ~ 500.00KVA CTRatio 10
Transf. Types  S7.59(11) Transf. Transf. No. 000001
Connection  YynO Operator No. Num_1

Lo/





Fig5  Setting screen

Secondary U：Inputting the working voltage before testing is the secondary rated voltage, which can input directing by using keyboard. Unit is KV.

Nominal Power: Rated capacity of testing transformer

Transf. Types ：Setting up the transformer model.  

Main setting model:  "SJ (73) Transf.", "S7.S9 (11) Transf. ", "S7.S9 (11) Power", "S13 Transf.", "Coating CRDT", "unCoating CRDT", "NonStand. Transf."

Connection : Testing model of connection of transformer , it can choose by using “←”、“→”according to nameplate

Testing Model:  including “3P4W No-Load””1Phase No-Load”, “3P3W No-Load”

PT Ratio: When testing voltage is over than the testing range of instrument, user can extend the voltage through connecting PT. This data is the ratio between the rated voltage and the connection PT(for example: 10KV/0.1KV PT =100). When it is not necessary to connect PT, the ratio is 1.

CT Ratio: same with Voltage ratio. When testing current is over than the testing range of instrument, user can extend the current through connecting CT. This data is the ratio between the rated current and the connection CT(for example: 10KV/0.1KV CT =100). When it is not necessary to connect CT, the ratio is 1.

Transf. No. : 6 digits of number in order to check and save.

Operator No.: Operator No. in order to check and save.

After inputting, Click “Save”, data can be saved in instrument in order to check in future; click “End”, cancel this operation, click “End” again, Screen returns to the part of setting up parameter.

4.2  3P4W No-LoadLossTesting 

[image: image7.png]3P4W No-LoadLossTesting

Ua: 0.00V
Ia: 0.00A
Pa: 0.00W

CNS Noload:
Tested T.Type: -

Data

Ub: 0.00V Uc: 0.00V

Ib: 0.00A Ic: 0.00A

Pb: 0.00W Pc: 0.00W
Noload Loss: 0.0W

NoLoad I(%): 0.0%

Result Exit





Fig6  3P4W No-LoadLossTesting
Changing the testing method to “3P4W No-Load” through moving the “↑”、“↓”、“←”、“→” on no-load test screen, and then beginning test(please see the following instruction)

Testing result shows Voltage( Ua、Ub、Uc); Current( Ia、Ib、Ic ) ; Active loss (Pa、Pb、Pc). Other result is Io ,Po after calibrating and the level No. of transformer according to the test. Click “Print “to print , click “save “ to save.
The instrument will automatically save the result when the input voltage is up to rated voltage under 3 phase no-load test. Upon saving , power can cut off. Data also can be saved through clicking “enter” when the input power is not up to rated voltage.

4.3 3P3W No-Load 
Changing the testing method to “3P3W No-Load” through moving the “↑”、“↓”、“←”、“→” on no-load test screen, and then beginning test(please see the following instruction)

[image: image8.png]3P3W No-LoadLossTesting

Uab: 0.00V
la: 0.0A

Pab: 0.00W

CNS Noload:
Tested T.Type: -

Data

Ucb: 0.00V
Ic: 0.00A

Pcb: 0.00W

Noload Loss:
NoLoad 1(%): 0.0%

Result Exit

0.0W





Fig7   3P3W No-Load

Testing result shows Voltage(Uab/Ucb); Current( Ia/Ic ) ; Active loss (Pab、Pcb). Other result is Io ,Po after calibrating and the level No. of transformer according to the test. Click “Print “to print , click “save “ to save.
The instrument will automatically save the result when the input voltage is up to rated voltage under test. Upon saving , power can cut off. Data also can be saved through clicking “enter” when the input power is not up to rated voltage.

4.4  1Phase No-Load Test

Changing the testing method to “1Phase No-Load Test” through moving the “↑”、“↓”、“←”、“→” on no-load test screen, and then beginning test(please see the following instruction)

[image: image9.png]1Phase No-Load Testing

CNS Noload:
Tested T.Type: -

Data

Ua: 0.00V
Ia: 0.00A

Pa: 0.00W

Noload Loss:
NoLoad 1(%): 0.0%

Exit

0.0W





 Fig8  1Phase No-Load Test

Testing result shows Voltage (Ua); Current ( Ia ) ; Active loss (Pa). Other result is Io, Po after calibrating and the level No. of transformer according to the test. Click “Print “to print, click “save “to save.
The instrument will automatically save the result when the input voltage is up to rated voltage under single phase no-load test. Upon saving , power can cut off. Data also can be saved through clicking “enter” when the input power is not up to rated voltage.

5 Load Loss

“Load Loss” and “No-Load Loss” have similar screens and operations. The difference as follows:

[image: image10.png]Transformer LoadLoss Testing Setup

Primary U 10.00KV
Testing Temp.  20.0°C
Nomial Power 500KVA

Calibrate Temp. 75.0°C

Connection  YynO

Testing Model
PT Ratio
CT Ratio

Transf. No.

Operator No.

3P4W Load

10
10
000001

Num_01

(oo





Figure 9 Parameter settings interface for the load test 

5.1 Parameter setting

It needs to set some necessary parameters before the load test. On the interface of setting load test,moving cursor by the “↑”,“↓”,“←”,“→”keys, and enter "parameter Settings" interface.(Figure 9).

Primary voltage: Inputting the working voltage of the transformer correctly before the determination of the transformer capacity, this is a rated voltage of the transformer, which can be inputted directly through the digital keyboard. Unit: KV.

Testing Temp.: It is used to test the load loss correction to the standard load test conditions(Such as 75 ℃),the correction formula of load loss is 
[image: image11.wmf]75

×

kk

PKP

=

,K is temperature coefficient of resistance .The algorithm is 
[image: image12.wmf](23575)

(235)

K

t

+

=

+

,T is the actual temperature when test. The temperature correction formula of impedance voltage is 
[image: image13.wmf]222

(/10)(1)

KTKTN

UkUPSK

=+-

,
[image: image14.wmf]KT

U

 is the actual percentage of impedance voltage and temperature,
[image: image15.wmf]KT

P

 is the actual load loss of the current temperature,
[image: image16.wmf]N

S

is the actual tested transformer rated capacity. Please use infrared thermometer to measure the current temperature of the transformer under test, and input through the digital key.

Nominal capacity: the participant rated capacity of transformer.

Calibrate Temp.: As mentioned in the "Testing Temp." item, load loss experiment results need to be checked, proofreading to uniform temperature range, this is the uniform temperature range. General calibration of oil-immersed transformer temperature is 75 ℃, while the dry type transformer has a variety of different temperature correction.

Connection: The test method of transformer connection group. Please accord to the nameplate to choose by press ,“←”,“→”button.

Testing Model :This project contains the content of four modes, they are "In3P4W Load", "3P4W Load", "1Phase Load", "3P3W Load”, the specific connection mode please refer to the state later.

PT Ratio: When the measured voltage value exceeds range after the instrument itself, users can extend the testing range by the external voltage transformer, This parameter is for outside enlarge voltage transformer ratio, (for example: 10 kV / 0.1 kV voltage transformer, should enter 100).When not use outside enlarge voltage transformer, please enter 1.

CT Ratio: same with Voltage ratio. When testing current is over than the testing range of instrument, user can extend the current through connecting CT. This data is the ratio between the rated current and the connection CT (for example: 10KV/0.1KV CT =100). When it is not necessary to connect CT, the ratio is 1.

Transf. No. : 6 digits of number in order to check and save.

Operator No.: Operator No. in order to check and save.

When you input the correct data, you can click the "ok" button to save the data. Click the "end" button, you can cancel the input. Again, click the "end" button to return to main interface.

5.2 In3P4W Load Loss Testing(Active Three-phase Load Testing)

“In3P4W Load” approach is to use an instrument built-in three-phase power supply as a test power. Results show Uk, that is the after correction impedance voltage, adjusted the load loss Pk (75).The correction here refers to the value of the load voltage and load loss which measured by the load test when not under the rated power, correct to the rated power, and at the meaning time, the temperature correct to the standard test conditions (such as 75 ℃).the calibration of load test are all this mean, so not repeat any more. Please see figure 10.

[image: image17.png]In3P4W LoadLossTesting

Ua: 0.00V Ub: 0.00V Uc: 0.00V

Ia: 0.00A Ib: 0.00A Ic: 0.00A
Pa: 0.00W Pb: 0.00W Pc: 0.00W
Pk(75): 0.000KW

Uk(75): 0.00KW

Data Result Exit





Figue 10 In3P4W LoadLossTesting
5.3 3P4W Load Loss Testing

The function is basically the same between the 3P4W Load Loss and the In3P4W Load Loss, the only difference is the selection way of test power source, Active three-phase load test is to use an instrument built-in three-phase power supply as a test power, while the three-phase load test is the external power supply as a test power. Detailed wiring method, please refer to the back of the detailed description.

Test and the content display is the same with the active three-phase load test, Not repeat description any more.

Note: The instrument will automatically latch the result if the short circuit current reaches rated current of the transformer when testing. The testing power supply can be stopped once the test result is latched. In the process of test, you can also latch test results temporarily by clicking the "ok" button.

5.4 The other load testing

Wiring method is the main difference between 3P3W Load Loss and 1Phase Load; Details please see the back of the specific wiring instructions.

The test results are the basic same. Refer to figure 11 and figure 12.
[image: image18.png]3P3W LoadLossTesting

Uab: 0.00V
la: 0.0A

Pab: 0.00W

Data

Ucb: 0.00V
Ic: 0.00A

Pcb: 0.00W

PK(75): 0.000KW

Uk(75): 0.00KW

Exit





Figure 11 3P3W Load Loss

[image: image19.png]1Phase Load Testing

Ua: 0.00V
Ia: 0.00A
Pa: 0.00W
Pk(75): 0.000KW
Uk(75): 0.00KW

Data Result Exit





Figure 12  1Phase Load Testing

Note: The instrument will automatically latch the result if the short circuit current reaches rated current of the transformer when testing. The testing power supply can be stopped once the test result is latched. In the process of test, you can also latch test results temporarily by clicking the "ok" button.

6. Data Browser

Selected in the main menu interface "Data Browser", and click "ok" button and enter the data query interface (FIG. 13, 14)

Data shown here is the saving data of test result.When the preservation of "no-load test" or "load test" data,The lower right corner of figure 13 will display“→”，then,By pressing the “←”、“→” button to view the capacity test, no-load test, load test of the test results.Using "up", "down" to see other transformers storage page.Click "print" button to print the current page data.Click "delete" button, you can delete the current record or all of the records. Click the "end", that is, back to the main menu.

[image: image20.png]Data Browser

Transf.No.| Load Loss
PrimaryU|  CNSLoadLoss
SecondaryU~ Load Loss(75°C)
Test Temp. T LoadLoss Error
Tapping Step ] Tested Uz%
Nominal Power Tested R
Nominal Uz%6 = Tested Type
Connection Final T.Power

Operator Tested T.Power





Figure 13  Store data browsing interface—1

[image: image21.png]Data Browser

4=

No-Load Loss :

No-Load I(%) :

CNS NoloadLoss :

Tested T.Type:

Load Test Result

Calibration Load Loss :

Short-Circuit impedance :





Figure  14 Store data browsing interface—2
 7.Data upload

When you need to upload data to the host computer database, please open the instrument power, and then through the data line to connect the instrument with the computer correctly, according to the prompts to install the driver. Operation capacity of communication software, select the corresponding COM port and click on the Toolbar "Upload" button, open upload interface , click the "upload data", at this time, can be stored in the instrument of the previous test data uploaded to the computer. For the use of communication software of the capacity, please refer to the operation instruction.

V.wiring method

1. Concept Introduction

No load test: Please apply rated AC rated voltage under the rated frequency (usually the 50Hz sine wave) from one winding of transformer; other windings need to open circuit. The measurement result includes the no-load current and no-load loss.

If the test condition is limited, the power supply voltage is not up to the rated voltage, which can be tested under the condition of non-rated voltage. But the error of the measurement is relatively large, which is generally only used to check the transformer. Only when the test voltage is up to 70% of the rated voltage, it can be measured no-load current and no-load loss more accurately.

Load test: Please apply rated AC rated voltage under the rated frequency (usually the 50Hz sine wave), this need to short circuit of one winding of transformer (generally the low voltage current side), and then access the Ac rated voltage from tap of coil (generally high voltage side), The current of testing terminal in winding need to up to the rated current, and then testing the load loss and load voltage

2. Wiring Method of Transformer Capacity Test and Active Load Test

As shown in figure 15. The “binding post “is just to facilitate your understanding. The actual wiring is a seven holes socket.

[image: image22.jpg]Instrument end connection (sketch map)

JoWIOFSURI]




Figure 15 Capacity test and the connection method of active load test

3. Wiring Method of Single-Phase Transformer No-load Loss with Single-phase Power

When the voltage and current of the test are within the range of the instrument's measuring range, please apply the current and voltage directly to the instrument as shown in Figure 16. When the no-load loss is tested, the general low voltage side is the tested side. High voltage side is non-tested part which need to open circuit.

When the test voltage exceeds the voltage range of the instrument, please use the voltage transformer, current transformer, according to figure 17.

   When the test current exceeds the test range of the instrument, but the voltage is not over, please use the current transformer to access the current, the voltage directly access, according to figure 18.

4. Wiring method of Single-phase Transformer Short-circuit Loss with Single-phase Power

It is similar with the wiring method of no-load loss test for single phase power and single phase transformer; please refer to figure 16, figure 17, figure 18. The difference is just that the short-circuit loss test, the general high voltage side for the tested side. The low pressure side is a non-tested end, in addition, the non-tested side needs to manually short circuit.

[image: image23.jpg]Transformer

2

S

Sutngle-phase
power.





      Figure 16 Direct access method for single phase load test

[image: image24.png]LR

Ta e
Uo

HEPE





   Figure 17 Single phase no-load test external PT and CT access method

[image: image25.png]transformer





Figure 18 Single phase no-load test external CT access method

5. Wiring Method of Three-phase Transformer No-load Loss with Single-phase power

When finding the loss is over than the standard value after three phases no-load loss test, operator needs to test three losses, compare and analyze the no-load loss, observe the distributed situation of no-load loss in every phase, in order to check whether has shortage in every winding resistance. Basic method is to take the three phase transformer as three pieces separated transformers to add voltage. That means one phase of transformer need to be short circuit and then apply the voltage on the other two phases to test the no-load loss and no-load current Figure 19.20.21 are the different wiring methods according to winding resistance of tested transformer  Figure 22, 23, 24 are the different methods according to the range of current and voltage.

[image: image26.png]b e b e 3 o
=i g
add voltage add voltage add voltage

on a,b phase on bc phase on c,a phase




Figure 19 schematic diagram of △ Pressurized winding 
[image: image40]
Figure 20 schematic diagram of neutral point  Y Pressurized winding
[image: image27.png]



Figure 21 schematic diagram of Pressurized winding without neutral point Y

Attention: When the wiring method is the Y, please apply double voltage of phase voltage on product, When the wiring method without Y, operator should make the secondary winding resting short circuit because we can make non-pressurized winding resistance short circuit without the neural point , 
[image: image28.jpg]Transformer

o 0 ©°
5 o @

inglo-phase powe:





Figure22 Direct access method of single-phase power measurement of three-phase transformer with no-load loss 

[image: image29.jpg]Transformer

© © o

o o o

e

.
] S
Single-phase pover





   Figure 23 The single-phase power measurement of three-phase transformer no-load loss external PT and CT access method.

[image: image30.jpg]Transformer

0 0o 0

o

5 1

Goe

o
2
mﬁ:ﬂ:_:n poer




      Figure 24 The single-phase power measurement of three-phase transformer no-load loss external CT access method

6. Wiring Method of Three-phase Transformer Short Circuit Loss with Single-phase Power.

Under the influence of power (no three phase power supply or power supply capacity is small), or fault phase need to be checked during the producing process or operation, single-phase power can applied to short-circuit loss test. The test method is to make short circuit of three phase low voltage of transformer, and then test high voltage side. “Single phase test” in short circuit test of this instrument can directly test three phases without exiting current program, and then show the total result. 
According to the winding connection mode of the tested transformer can be divided into two cases: A，when the pressurized winding is type △ , please check figure 19. The difference is that the non pressurized side (usually for the low voltage side) of the three-phase outlet side need to short circuit manually. The current in the windings should be2/√3 times than the rated current of the transformer. B. When pressure winding is Y - type, the connection of the inflating is like picture 21, the difference is that the pressure side of the three-phase side need to be short circuit manually.  
 According to the measured voltage, current and the voltage of instrument and current measuring range, there are also three cases which is similar with the three types of no-load loss three phase transformer under single phase power, please refer to figure 22, 23, 24 shows the connection mode, difference is, the secondary side should short circuit.

7、Wiring method and two elements No-load Loss Test (three-phase three wire power) 

Please let short circuit of non-tested terminal and refer to figure 25 when the voltage and current does not exceed the test range of the instrument, please refer to figure 25; when the voltage exceeds the test range of the instrument, please refer to figure 26 shows connection; when the test current exceeds the instrument testing scope and voltage does not exceed the test range of the instrument, please refer to figure 27. When the no-load loss is tested, the general low voltage side is the tested side. High voltage side is non-tested end that need to open circuit

   Note: this method is equivalent to the two power meter test method, only measuring Uab and Ubc phase voltage value result, the average value of the two; at the same time also only no-load loss measurement of Pab and Pcb phase loss, total loss is the sum of two phases.

[image: image31.jpg]Transformert

Une-

(L

+Ic-

R
s 8 8
hree-phase pover

T





Figure25   Two components no-load loss test method of direct wiring diagram
[image: image32.jpg]Transformer

o0 o
o o o

£8
_ ==
PAES

o

v
§ o
o e o o
2 e 8 8
Three-phase power





Figure26 No-load loss test method of two components of PT and CT wiring diagram
[image: image33.jpg]Transformer

© 0 o
o

36 6 e
S 8 8 5

Three-phase power





Figure27 Two components no-load loss test method of CT wiring diagram

8、Wiring method and two elements Load Loss Test (three-phase three wire power) 

Basically similar with wiring mode of two element method of no-load loss can refer to figure 25, 26, and 27. The difference is the high-voltage side to test end.  Low voltage side for the test side when test short circuit loss. In addition, the test need artificial nipple.  Such as high or medium voltage side outlet casing containing the ring current transformer, prior to the test will be a good short circuit of current transformer secondary side.

9、Wiring method and Three phase No-load loss Test (three-phase four wire power)
Please let short circuit of non-tested terminal and refer to figure 25 when the voltage and current does not exceed the test range of the instrument, please refer to figure 28; when the voltage exceeds the test range of the instrument, please refer to figure 29 shows connection; when the test current exceeds the instrument testing scope and voltage does not exceed the test range of the instrument, please refer to figure 30. When the no-load loss is tested, the general low voltage side is the tested side. High voltage side is non-tested end that need to open circuit.

10、Wiring method and Three phase Load loss Test (three-phase four wire power)
Basically similar with wiring mode of two element method of no-load loss can refer to figure 28, 29, and 30. The difference is the high-voltage side to test end.  Low voltage side for the test side when test no-load loss. In addition, the test need artificial nipple.  Such as high or medium voltage side outlet casing containing the ring current transformer, prior to the test will be a good short circuit of current transformer secondary side.

[image: image34.jpg]Transformer

© 0 ©

“né

né [

ole|e
S(8]8
sy
S &2 8
i
o ¢

4
6 ¢ Lo

o
S 8 8 8
Three-phase power





Figure28     Direct wiring diagram of three-phase no-load loss test
[image: image35.jpg]Transformer
o0 o0

TR
[ C

va
w e

X

R

O+Ia-
+1Ib-

e +Ic-o [Uce

=
S

[ [
¢ 4 o
8 38

Three-phaswe power)





Figure29 Three-phase no-load loss test of PT and CT wiring diagram
[image: image36.png]isisis

Transformer




Figure29 Three-phase no-load loss test of CT wiring diagram
VI. Note

1、The connecting method of testing line should operate according to instruction, otherwise it will affect the result

2、 The wiring should be proceeded when the testing line has grounded in order to electric shock of induced voltage.

3、Please ensure the voltage ,current under the testing range. It need to connect with CT PT to expand the range , or testing result will not right.

4、When instrument do capacity test, active load test, load test, 

Please ensure the short circuit of the no-pressure side, otherwise it will affect the result.

When doing the load test, if there is the power casing of HV LV side(Eg: CT), Please ensure short circuit of secondary winding of CT.

VII.Packing List
	No.
	Item
	Qty

	1
	Test line
	1

	2
	short wiring
	1

	3
	Charger
	2

	4
	Ground cable
	1

	5
	Print paper
	1


Appendix-Catalogue

A

   The following is the rule of performance level code of 10KV,35KV oil transformer and the partial no-load loss of 10KV , 35KV, according to the JB/T3837-1996、GB/T 6451—1999和GB/T 10228—1997 and so on.

A1 the Provision of Performance Level of Three-phase Oil Power Transformer

	performance level code
	Voltage class (kV)
	Parameter

	
	
	no-load loss
	

	7
	6、10
	GB/T 6451 GroupII
	GB/T 6451

	
	≥35
	GB/T 6451
	

	8
	6、10
	GB/T 6451 GroupI
	

	
	≥35
	Less than GB/T 6451 10% on average
	

	9
	6、10
	Distribution TransformerA_2

	
	6、10
	Power transformer is less than GB/T 6451 Group I 10% on average
	less thanGB/T 6451 on average10%

	
	≥35
	Less than GB/T 6451 20% on average
	

	10
	6、10
	GB/T 6451 Group I 20% on average
	less than GB/T 6451 10% on average

	
	≥35
	less than GB/T 6451 30% on average
	

	11
	6、10
	less than GB/T 6451 Group I 30% on average
	

	
	≥35
	less than GB/T 645140% on average
	


A2 S9-30~1600/6、10 Parameters of Distribution Transformer

	Rated capacity kVA
	No-load loss W
	Load loss W

	30
	130
	600

	50
	170
	870

	63
	200
	1040

	80
	250
	1250

	100
	290
	1500

	125
	340
	1800

	160
	400
	2200

	200
	480
	2600

	250
	560
	3050

	315
	670
	3650

	400
	800
	4300

	500
	960
	5100

	630
	1200
	6200

	800
	1400
	7500

	1000
	1700
	10300

	1250
	1950
	12800

	1600
	2400
	14500


A3 10kV Oil Testing Transformer No-load Loss with Different Levels

	Transformer capacity
	S7
	S8
	S9power transformer
	S9distribution transformer
	S10
	S11

	30
	0.15
	0.14
	0.126
	0.13
	0.112
	0.098

	50
	0.19
	0.19
	0.171
	0.17
	0.152
	0.133

	63
	0.22
	0.22
	0.198
	0.2
	0.176
	0.154

	80
	0.27
	0.25
	0.225
	0.25
	0.2
	0.175

	100
	0.32
	0.29
	0.261
	0.29
	0.232
	0.203

	125
	0.37
	0.34
	0.306
	0.34
	0.272
	0.238

	160
	0.46
	0.39
	0.351
	0.4
	0.312
	0.273

	200
	0.54
	0.47
	0.423
	0.48
	0.376
	0.329

	250
	0.64
	0.57
	0.513
	0.56
	0.456
	0.399

	315
	0.76
	0.68
	0.612
	0.67
	0.544
	0.476

	400
	0.92
	0.81
	0.729
	0.8
	0.648
	0.567

	500
	1.08
	0.97
	0.873
	0.96
	0.776
	0.679

	630
	1.3
	1.15
	1.035
	1.2
	0.92
	0.805

	800
	1.54
	1.4
	1.26
	1.4
	1.12
	0.98

	1000
	1.8
	1.65
	1.485
	1.7
	1.32
	1.155

	1250
	2.2
	1.95
	1.755
	1.95
	1.56
	1.365

	1600
	2.65
	2.35
	2.115
	2.4
	1.88
	1.645


A4 35kV Oil Testing Transformer No-load Loss with Different Levels

	Transformer capacity
	S7
	S8
	S9
	S10
	S11

	50
	0.27
	0.24
	0.216
	0.189
	0.162

	100
	0.37
	0.34
	0.296
	0.259
	0.222

	125
	0.42
	0.38
	0.336
	0.294
	0.252

	160
	0.47
	0.41
	0.376
	0.329
	0.282

	200
	0.55
	0.48
	0.44
	0.385
	0.33

	250
	0.64
	0.57
	0.512
	0.448
	0.384

	315
	0.76
	0.68
	0.608
	0.532
	0.456

	400
	0.92
	0.82
	0.736
	0.644
	0.552

	500
	1.08
	0.97
	0.864
	0.756
	0.648

	630
	1.3
	1.16
	1.04
	0.91
	0.78

	800
	1.54
	1.39
	1.232
	1.078
	0.924

	1000
	1.8
	1.65
	1.44
	1.26
	1.08

	1250
	2.2
	1.96
	1.76
	1.54
	1.32

	1600
	2.65
	2.37
	2.12
	1.855
	1.59

	2000
	3.4
	2.9
	2.72
	2.38
	2.04

	2500
	4
	3.5
	3.2
	2.8
	2.4

	3150
	4.75
	4.3
	3.8
	3.325
	2.85

	4000
	5.65
	5.15
	4.52
	3.955
	3.39

	5000
	6.75
	6.1
	5.4
	4.725
	4.05

	6300
	8.2
	7.3
	6.56
	5.74
	4.92

	8000
	11.5
	10
	9.2
	8.05
	6.9

	10000
	13.6
	11.8
	10.88
	9.52
	8.16

	12500
	16
	14
	12.8
	11.2
	9.6

	16000
	19
	17
	15.2
	13.3
	11.4

	20000
	22.5
	20.1
	18
	15.75
	13.5

	25000
	26.6
	23.9
	21.28
	18.62
	15.96

	31500
	31.6
	28.5
	25.28
	22.12
	18.96


A5 10KV,35KV Dry Type Testing Transformer No-load Loss with Different Levels


	Transformer capacity
	10k(SC)dry type transformer
	10kV (SG)dry type transformer
	35kV(SC) dry type transformer

	
	S7
	S8
	S9
	S10
	S7
	S8
	S9
	S10
	S8
	S9
	S10

	30
	0.27
	0.24
	0.216
	0.192
	0.31
	0.28
	0.252
	0.224
	0.4
	0.36
	0.32

	50
	0.37
	0.34
	0.306
	0.272
	0.39
	0.36
	0.324
	0.288
	0.5
	0.45
	0.4

	80
	0.5
	0.46
	0.414
	0.368
	0.49
	0.46
	0.414
	0.368
	0.578
	0.52
	0.462

	100
	0.59
	0.5
	0.45
	0.4
	0.54
	0.51
	0.459
	0.408
	0.733
	0.66
	0.586

	125
	0.69
	0.59
	0.531
	0.472
	0.63
	0.6
	0.54
	0.48
	0.844
	0.76
	0.675

	160
	0.82
	0.68
	0.612
	0.544
	0.74
	0.7
	0.63
	0.56
	0.956
	0.86
	0.765

	200
	0.93
	0.78
	0.702
	0.624
	0.86
	0.82
	0.738
	0.656
	1.067
	0.96
	0.854

	250
	1.1
	0.9
	0.81
	0.72
	1
	0.95
	0.855
	0.76
	1.2
	1.08
	0.96

	315
	1.3
	1.1
	0.99
	0.88
	1.2
	1.1
	0.99
	0.88
	1.45
	1.3
	1.16

	400
	1.59
	1.22
	1.098
	0.976
	1.4
	1.3
	1.17
	1.04
	1.7
	1.53
	1.36

	500
	1.85
	1.45
	1.305
	1.16
	1.65
	1.5
	1.35
	1.2
	2
	1.8
	1.6

	630
	2.1
	1.68
	1.512
	1.344
	1.9
	1.75
	1.575
	1.4
	2.3
	2.07
	1.84

	800
	2.3
	1.9
	1.71
	1.52
	2.35
	2.12
	1.908
	1.696
	2.7
	2.43
	2.16

	1000
	2.65
	2.21
	1.989
	1.768
	2.75
	2.48
	2.232
	1.984
	3
	2.7
	2.4

	1250
	3.28
	2.61
	2.349
	2.088
	3.3
	2.98
	2.682
	2.384
	3.5
	3.15
	2.8

	1600
	3.75
	3.06
	2.754
	2.448
	3.8
	3.42
	3.078
	2.736
	4
	3.6
	3.2

	2000
	4.7
	4.15
	3.735
	3.32
	4.6
	4.15
	3.735
	3.32
	4.7
	4.23
	3.76

	2500
	5.7
	5
	4.5
	4
	5.54
	5
	4.5
	4
	5.5
	4.95
	4.4

	3150
	——
	——
	——
	——
	——
	——
	——
	——
	6.5
	5.85
	5.2

	4000
	——
	——
	——
	——
	——
	——
	——
	——
	7.5
	6.75
	6


B  GB/T 6451 and GB/T 10228 Appendix B GB/T 6451 and GB/T 10228
B1 10KV Double Winding Transformer with Off-circuit Tap Changing

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	No-load loss kW
	Load loss                 kW
	No-load  current   %
	Short circuit impedance %

	
	HV   kV
	Tapping range of HV %
	LV  kV
	
	
	
	
	

	
	
	
	
	
	
	Yyn0
	Dyn11/Yzn11
	Yyn0
	Dyn11/Yzn11
	

	30
	6     6.3    10   10.5   11
	±5
	0.4
	Yyn0
	0.14
	0.80
	0.80
	2.80
	2.80
	4.00

	50
	
	
	
	Yyn0 Dyn11 Yzn11
	0.19
	1.15
	1.25
	2.50
	2.70
	

	63
	
	
	
	
	0.22
	1.40
	1.50
	2.40
	2.60
	

	80
	
	
	
	
	0.25
	1.65
	1.80
	2.20
	2.40
	

	100
	
	
	
	
	0.29
	2.00
	2.15
	2.10
	2.30
	

	125
	
	
	
	
	0.34
	2.45
	2.55
	2.00
	2.20
	

	160
	
	
	
	
	0.39
	2.85
	3.10
	1.90
	2.10
	

	200
	
	
	
	
	0.47
	3.50
	3.60
	1.80
	2.00
	

	250
	
	
	
	
	0.57
	4.00
	4.10
	1.70
	1.90
	

	315
	
	
	
	
	0.68
	4.80
	4.90
	1.60
	1.80
	

	400
	
	
	
	
	0.81
	5.80
	6.00
	1.50
	1.70
	

	500
	
	
	
	
	0.97
	6.90
	7.15
	1.40
	1.60
	

	630
	
	
	
	Yyn0 Dyn11 
	1.15
	8.10
	8.50
	1.30
	1.50
	4.50

	800
	
	
	
	
	1.40
	9.90
	10.40
	1.20
	1.40
	

	1000
	
	
	
	
	1.65
	11.60
	12.20
	1.10
	1.30
	

	1250
	
	
	
	
	1.95
	13.80
	14.50
	1.00
	1.20
	

	1600
	
	
	
	
	2.35
	16.50
	17.30
	0.90
	1.10
	


B2  35KV Double Winding Transformer with Off-circuit Tap Changing

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	No-load loss kW
	Load loss                 kW
	No-load current     %
	Short circuit impedance %

	
	HV   kV
	Tapping range of HV %
	LV   kV
	
	
	
	
	

	
	
	
	
	
	
	Group I
	Group II
	Group I
	Group II
	

	50
	35
	±5
	0.4
	Yyn0
	1.35
	0.24
	0.27
	2.00
	2.80
	6.50

	100
	
	
	
	
	2.25
	0.34
	0.37
	1.80
	2.60
	

	125
	
	
	
	
	2.65
	0.38
	0.42
	1.75
	2.50
	

	160
	
	
	
	
	3.15
	0.41
	0.47
	1.65
	2.40
	

	200
	
	
	
	
	3.70
	0.48
	0.55
	1.55
	2.20
	

	250
	
	
	
	
	4.40
	0.57
	0.64
	1.40
	2.00
	

	315
	
	
	
	
	5.30
	0.68
	0.76
	1.40
	2.00
	

	400
	
	
	
	
	6.40
	0.82
	0.92
	1.30
	1.90
	

	500
	
	
	
	
	7.70
	0.97
	1.08
	1.30
	1.90
	

	630
	
	
	
	
	9.20
	1.16
	1.30
	1.25
	1.80
	

	800
	
	
	
	
	11.00
	1.39
	1.54
	1.05
	1.50
	

	1000
	
	
	
	
	13.50
	1.65
	1.80
	1.00
	1.40
	

	1250
	
	
	
	
	16.30
	1.96
	2.20
	0.85
	1.20
	

	1600
	
	
	
	
	19.50
	2.37
	2.65
	0.75
	1.10
	


B3  Double Winding Regulating Transformer with Off-circuit Tap Changing

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	Load loss kW
	No-load loss                kW
	No-load current    %
	Short circuit impedance %

	
	HV   kV
	Tapping range of HV %
	LV  kV
	
	
	
	
	

	
	
	
	
	
	
	Group I
	Group II
	Group I
	Group II
	

	800
	35
	±5
	3.15  6.3  10.5
	Yd11
	11.00
	1.39
	1.54
	1.05
	1.50
	6.50

	1000
	
	
	
	
	13.50
	1.65
	1.80
	1.00
	1.40
	

	1250
	
	
	
	
	16.30
	1.96
	2.20
	0.90
	1.30
	

	1600
	
	
	
	
	19.50
	2.37
	2.65
	0.85
	1.20
	

	2000
	
	
	
	
	21.50
	2.90
	3.40
	0.75
	1.10
	

	2500
	
	
	
	
	23.00
	3.50
	4.00
	0.75
	1.10
	

	3150
	35   38.5
	±5
	3.15  6.3  10.5
	Yd11
	27.00
	4.30
	4.75
	0.70
	1.00
	7.00

	4000
	
	
	
	
	32.00
	5.15
	5.65
	0.70
	1.00
	

	5000
	
	
	
	
	36.70
	6.10
	6.75
	0.60
	0.90
	

	6300
	
	
	
	
	41.00
	7.30
	8.20
	0.60
	0.90
	7.50

	8000
	
	±2×2.5
	3.15  3.3   6.3   6.6  10.5   11
	YNd11
	45.00
	10.00
	11.50
	0.55
	0.80
	

	10000
	
	
	
	
	53.00
	11.80
	13.60
	0.55
	0.80
	

	12500
	
	
	
	
	63.00
	14.00
	16.00
	0.50
	0.70
	8

	16000
	
	
	
	
	77.00
	17.00
	19.00
	0.50
	0.70
	

	20000
	
	
	
	
	93.00
	20.10
	22.50
	0.50
	0.70
	

	25000
	
	
	
	
	110.00
	23.90
	26.60
	0.40
	0.60
	

	31500
	
	
	
	
	132.00
	28.50
	31.60
	0.40
	0.60
	


B4  10KV Dry Type Distribution Transformer without Encapsulated Coil and Excitation

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	Load loss under different insulation endurance class W
	No-load loss                 W
	No-load current         %
	Short circuit impedance %

	
	HV   kV
	Tapping range of HV %
	LV kV
	
	B(100℃)
	F(120℃)
	H(145℃)
	Group I
	Group II
	Group I
	Group II
	

	30
	6     6.3   6.6    10      10.5   11
	±5 or  ±2×2.5
	0.4
	Yyn0  Dyn11
	760
	810
	980
	280
	310
	3.20
	3.50
	4.00

	50
	
	
	
	
	1100
	1170
	1490
	360
	390
	2.80
	3.00
	

	80
	
	
	
	
	1590
	1680
	2150
	460
	490
	2.60
	2.80
	

	100
	
	
	
	
	1850
	1950
	2550
	510
	540
	2.40
	2.60
	

	125
	
	
	
	
	2150
	2280
	3050
	600
	630
	2.20
	2.40
	

	160
	
	
	
	
	2500
	2650
	3650
	700
	740
	2.20
	2.40
	

	200
	
	
	
	
	2950
	3130
	4680
	820
	860
	2.00
	2.20
	

	250
	
	
	
	
	3450
	3650
	5500
	950
	1000
	2.00
	2.20
	

	315
	
	
	
	
	4100
	4340
	6600
	1100
	1200
	1.80
	2.00
	

	400
	
	
	
	
	4950
	5250
	7800
	1300
	1400
	1.80
	2.00
	

	500
	
	
	
	
	6000
	6360
	9350
	1500
	1650
	1.80
	2.00
	

	630
	
	
	
	
	7100
	7500
	10900
	1750
	1900
	1.60
	1.80
	

	630
	
	
	
	
	8050
	8550
	11500
	1680
	1830
	1.60
	1.80
	6

	800
	
	
	
	
	9700
	10300
	13600
	2120
	2350
	1.60
	1.80
	

	1000
	
	
	
	
	11600
	12300
	15700
	2480
	2750
	1.40
	1.60
	

	1250
	
	
	
	
	13900
	14700
	18400
	2980
	3300
	1.40
	1.60
	

	1600
	
	
	
	
	16700
	17700
	21300
	3420
	3800
	1.40
	1.60
	

	2000
	
	
	
	
	20000
	21200
	25000
	4150
	4600
	1.20
	1.40
	

	2500
	
	
	
	
	24500
	26000
	29100
	5000
	5540
	1.20
	1.40
	


B5  10KV Dry Type Power Transformer without Encapsulated Coil and Excitation


	Rated capacity kVA
	Voltage and range of tap
	Transformer connection
	Load loss under different insulation endurance class W
	No-load loss W
	No-load current %
	Voltage and range of tap %

	
	HV kV
	Tapping range of HV %
	LV kV
	
	B(100℃)
	F(120℃)
	H(145℃)
	
	
	

	630
	6     6.3   6.6     10   10.5   11
	±5 or  ±2×2.5
	3    3.15    6     6.3
	Yyn0  Dyn11
	8600
	9100
	12300
	1850
	1.80
	6.00

	800
	
	
	
	
	10300
	10900
	14500
	2250
	1.80
	

	1000
	
	
	
	
	12400
	13100
	16800
	2650
	1.80
	

	1250
	
	
	
	
	14800
	15700
	19700
	3180
	1.60
	

	1600
	
	
	
	
	17800
	18900
	22700
	3650
	1.60
	

	2000
	
	
	
	
	21400
	22600
	26700
	4400
	1.40
	

	2500
	
	
	
	
	26200
	27800
	3100
	5350
	1.40
	

	3150
	
	
	
	
	31500
	33400
	35700
	6420
	1.20
	7.00

	4000
	10    10.5    11
	
	
	
	38400
	40700
	43600
	7470
	1.20
	

	5000
	
	
	
	
	46500
	49200
	52700
	8750
	1.00
	

	6300
	
	
	
	
	56700
	60100
	64300
	10400
	1.00
	


B6 10kV Double Winding Regulating Transformer without Excitation

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	No-load loss kW
	Load loss kW
	No-load current%
	Short circuit impedance%

	
	HV kV
	Tapping range of HV%
	LV kV
	
	
	
	
	

	630
	6     6.3    10   10.5   11
	±5
	3    3.15   6.3
	Yd11
	1.15
	8.10
	13.00
	4.5

	800
	
	
	
	
	1.40
	9.90
	1.20
	5.5

	1000
	
	
	
	
	1.65
	11.60
	1.10
	

	1250
	
	
	
	
	1.95
	13.80
	1.00
	

	1600
	
	
	
	
	2.35
	16.50
	0.90
	

	2000
	
	
	
	
	2.80
	19.80
	0.90
	

	2500
	
	
	
	
	3.30
	23.00
	0.80
	

	3150
	
	
	
	
	3.90
	27.00
	0.80
	

	4000
	10   10.5   11
	
	3.15   6.3
	
	4.80
	32.00
	0.70
	

	5000
	
	
	
	
	5.70
	36.70
	0.70
	

	6300
	
	
	
	
	6.80
	41.00
	0.60
	

	
	
	
	
	
	
	
	
	


B7 10KV Encapsulated Coil Dry Type Distribution Transformer without Excitation

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	Load loss under different insulation endurance class W
	No-load loss                  W
	No-load current                 %
	Short circuit impedance  %

	
	HV   kV
	Tapping range of HV%
	LV   kV
	
	B(100℃)
	F(120℃)
	H(145℃)
	Group I
	Group II
	Group I
	Group II
	

	30
	6     6.3   6.6    10      10.5   11
	±5or  ±2×2.5
	0.4
	Yyn0  Dyn11
	780
	830
	890
	240
	270
	3.20
	3.50
	4.00

	50
	
	
	
	
	1100
	1170
	1260
	340
	370
	2.80
	3.00
	

	80
	
	
	
	
	1520
	1620
	1740
	460
	500
	2.60
	2.80
	

	100
	
	
	
	
	1740
	1850
	1990
	500
	590
	2.40
	2.80
	

	125
	
	
	
	
	2040
	2170
	2330
	590
	690
	2.20
	2.50
	

	160
	
	
	
	
	2350
	2500
	2680
	680
	820
	2.20
	2.50
	

	200
	
	
	
	
	2790
	2970
	3180
	780
	930
	2.00
	2.20
	

	250
	
	
	
	
	3050
	3240
	3480
	900
	1100
	2.00
	2.20
	

	315
	
	
	
	
	3840
	4080
	4380
	1100
	1300
	1.80
	2.00
	

	400
	
	
	
	
	4410
	4690
	5030
	1220
	1590
	1.80
	2.00
	

	500
	
	
	
	
	5400
	5740
	6150
	1450
	1850
	1.80
	2.00
	

	630
	
	
	
	
	6500
	6910
	7400
	1680
	2100
	1.60
	1.80
	

	630
	
	
	
	
	6600
	7010
	7520
	1620
	2000
	1.60
	1.80
	6

	800
	
	
	
	
	7700
	8180
	8770
	1900
	2300
	1.60
	1.80
	

	1000
	
	
	
	
	9000
	9560
	10300
	2210
	2650
	1.40
	1.60
	

	1250
	
	
	
	
	10700
	11400
	12200
	2610
	3280
	1.40
	1.60
	

	1600
	
	
	
	
	13000
	13800
	14800
	3060
	3750
	1.40
	1.60
	

	2000
	
	
	
	
	16000
	17000
	18300
	4150
	4700
	1.20
	1.40
	

	2500
	
	
	
	
	19000
	20200
	21700
	500
	5700
	1.20
	1.40
	


B8 10KV Encapsulated Coil Dry Type Power Transformer without Excitation

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	Load loss under different insulation endurance class W
	No-load loss W
	No-load current%
	Short circuit impedance%

	
	HV kV
	Tapping range of HV%
	LV kV
	
	B(100℃)
	F(120℃)
	H(145℃)
	
	
	

	630
	6     6.3   6.6     10   10.5   11
	±5or  ±2×2.5
	3    3.15    6     6.3
	Yd11
	7000
	7440
	7970
	1750
	1.80
	6.00

	800
	
	
	
	
	8200
	8710
	9340
	2000
	1.80
	

	1000
	
	
	
	
	9700
	10300
	11100
	2400
	1.80
	

	1250
	
	
	
	
	11500
	12300
	13100
	2800
	1.60
	

	1600
	
	
	
	
	14000
	14900
	16000
	3300
	1.60
	

	2000
	
	
	
	
	16700
	17800
	19100
	4500
	1.40
	

	2500
	
	
	
	
	19700
	21000
	22500
	5300
	1.40
	

	3150
	
	
	
	
	23000
	24500
	26200
	6300
	1.20
	

	4000
	10    10.5    11
	
	
	
	27700
	29500
	31600
	7500
	1.20
	7.00

	5000
	
	
	
	
	32800
	34900
	37400
	8900
	1.00
	

	6300
	
	
	
	
	39000
	41400
	44400
	10500
	1.00
	


B9 35KV Encapsulated Coil Dry Type Distribution Transformer without Excitation

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	Load loss under different insulation endurance class W
	No-load loss W
	No-load current%
	Short circuit impedance%

	
	HV kV
	Tapping range of HV%
	LV kV
	
	B(100℃)
	F(120℃)
	H(145℃)
	
	
	

	315
	35    38.5
	±5or  ±2×2.5
	0.4
	Yyn0  Dyn11
	4620
	4900
	5240
	1450
	2.00
	6.00

	400
	
	
	
	
	5950
	6300
	6750
	1700
	2.00
	

	500
	
	
	
	
	7300
	7750
	8320
	2000
	2.00
	

	630
	
	
	
	
	8500
	9030
	9680
	2300
	1.80
	

	800
	
	
	
	
	10000
	10700
	11400
	2700
	1.80
	

	1000
	
	
	
	
	11500
	12300
	13100
	3000
	1.80
	

	1250
	
	
	
	
	14000
	14900
	16000
	3500
	1.60
	

	1600
	
	
	
	
	17000
	18100
	19400
	4000
	1.60
	

	2000
	
	
	
	
	20000
	21300
	22800
	4700
	1.40
	

	2500
	
	
	
	
	24000
	25500
	27400
	5500
	1.40
	


B10 35KV Encapsulated Coil Dry Type Power Transformer without Excitation

	Rated capacity kVA
	Voltage and range of tap
	Transf. connection
	Load loss under different insulation endurance class W
	No-load loss W
	No-load current%
	Short circuit impedance%

	
	HV kV
	Tapping range of HV%
	LV kV
	
	B(100℃)
	F(120℃)
	H(145℃)
	
	
	

	800
	35    38.5
	±5or  ±2×2.5
	3.15  3.3     6     6.3    10   10.5   11
	Yyn0  Yd11
	10300
	11000
	11800
	2780
	1.90
	6.00

	1000
	
	
	
	
	12000
	12800
	13700
	3300
	1.90
	

	1250
	
	
	
	
	14200
	15100
	16200
	3870
	1.70
	

	1600
	
	
	
	
	17000
	18100
	19400
	4560
	1.70
	

	2000
	
	
	
	
	20000
	21300
	22800
	5250
	1.50
	7.00

	2500
	
	
	
	
	24000
	25500
	27400
	6000
	1.50
	

	31500
	
	
	
	
	27000
	28700
	30800
	7500
	1.30
	8.00

	4000
	
	
	
	
	32500
	34500
	37000
	8700
	1.30
	

	5000
	
	
	
	
	38500
	40900
	43900
	10400
	1.10
	

	6300
	
	
	
	
	45000
	47800
	51300
	12300
	1.10
	

	8000
	
	
	
	Yd11 YNd11
	50000
	53000
	56700
	14000
	1.00
	9.00

	10000
	
	
	
	
	61300
	65000
	69500
	16000
	1.00
	


� EMBED  \* MERGEFORMAT ���





� EMBED  \* MERGEFORMAT ���









[image: image41][image: image42.png]Huazheng



[image: image43.wmf]100

´

=

ref

Z

Z

z

[image: image44.wmf]r

S

U

Z

ref

2

=

[image: image45.png]_1234567893.unknown

_1234567895.unknown

_1234567896.unknown

_1234567897.unknown

_1234567894.unknown

_1234567891.unknown

_1234567892.unknown

_1234567890.unknown

