[image: image12.png]R e e N Huazheng




[image: image23.png]




HZJQ-3 

Transformer Oil BDV Tester

[image: image1.png]RO e TS

S — e —

Mﬂs-wkw%





Huazheng Electric Manufacturing (Baoding) Co., Ltd
Dear user:

Thank you for choosing HZJQ-3 Transformer Oil BDV Tester.

     We hope that this instrument can make your work easier and more enjoyable, so that you can get the feeling of office automation in the test and analysis work.

  Before using the instrument, please read this manual, and operate and maintain the instrument according to the manual to prolong its service life.

"Just a light press, the test will be completed automatically" is the operating characteristics of this instrument.

     If you are satisfied with this instrument, please tell your colleagues; if you are not satisfied with this instrument, please call (0312) 6775656 to tell you to serve you at all times-Baoding Huazheng Electric Manufacturing Co., Ltd., our company will definitely make you satisfied !

Contents

1I. Overview


2II. Key Function And Feature


2III. Major Technical Indicators


3IV.Operational Approach


8V.Precautions


8VI.Maintenance


9VII.Oil Cup Cleaning Method And Common Fault Clearances


10VIII.Packing List




I. Overview
Many power systems, railway systems, large-scale petrochemical plants and enterprises have a lot of electrical equipment whose internal insulation are mostly oil-filled insulation type, and therefore, test on insulating oil dielectric strength is common and necessary. To meet the needs of the market, we have developed and produced a series of insulating oil dielectric strength testers according to national standard GB/T507-2002, industry standard DL429.9-91 and the latest Electric Power Industry Standard DL/T846.7-2004 by ourselves. This instrument, by using a single-chip microcomputer as the core, can operate in full automation with high accuracy, greatly improving work efficiency and reducing the labor intensity of workers.

II. Key Function And Feature
1. With a microprocessor, three cups of a body, automatically fulfill the withstand voltage test for oil circulation with a range of 0～80KV (including boosting, maintaining, mixing, standing, calculation, printing and other operations) 
2. Display by large-screen LCD and prompt by English menu.
3.The instrument has the advantages of simple operation, the operators only need a simple setup, instrument will automatically set in accordance with 1-3 oil sample pressure test. Each oil sample, each breakdown voltage value and the number of cycles will be automatically stored,after completion of the test, the thermal printer can print out the breakdown voltage value and average value.

4.Maintain power down, can store 100 experimental results, and can display the current environment temperature and humidity.

5.Using single chip microcomputer control speed boost, voltage frequency accuracy to 50HZ, makes the whole process easy to control.

6. Equipped with over-voltage, over-current and limit protections to ensure the safety of operators
7. With the function of displaying the measured temperature and system clock.
8. Communicate with computer with a standard RS232 interface.

III. Major Technical Indicators 
1、Output voltage: 0～80KV (optional)  

2、Voltage distortion rate: ＜3%

3、Voltage raising speed:0.5～5 KV/S (adjustable)

4、Standing time: 15min (adjustable)
5、Boosting interval: 5 min (adjustable)
6、Boosting times: 0～10 time(s) 
7、Booster capacity: 1.5kVA x 2
8、Measuring accuracy: ±3％
9、Supply voltage: AC220V±10%  50Hz±1 Hz
10、Power: 200W

11、Applicable temperature: 0℃～45℃

12、Applicable humidity: ＜75%RH

Overall dimension: 585*390*410mm
IV.Operational Approach

1．Preparation Before Test

①Connect the earth terminal (on the right side of the equipment) to the earth wire firmly before start the equipment.

②Sample the oil according to relevant standard. Adjust the electrode distance inside the oil cup with standard gauge. Clean the cup according to relevant requirements. Pour the sample into the cup and close the cap.

③ Switch on AC220V power supply after the above items are confirmed, ready for the test.

2．Start Testing

① Press the power switch and then enter the following interface:

[image: image2.png]Oil Breakdown Tester





Fig. 1 Main Interface

② Setting of system parameters:
Press “Settings” key and enter the following interface:

[image: image3.png]Q’: Parameter Settings

Voltage Settings
Max.kV: 80kV + —
Stirrer: Continuous| ;-
Stand Time: 15 min + -—
Rest Time: 5 min + -—
Test Count:  10times| .} ]
Volt.Rate: 2. 5kV/s| 0 L]

Print Settings
Auto Print: Yes + —

e — —
Time Settings

12011-08-09 10:33:15
R — —
Cups Selection

Cupl: Yes|





Fig.2 Parameter Set

Setting of voltage boosting: the user may select based on actual demand.

 SHAPE  \* MERGEFORMAT 



Cup bit setting: according to the user testing oil sample, placed oil cup bit shall prevail.

 SHAPE  \* MERGEFORMAT 



Print settings: Users choose for themselves.

 SHAPE  \* MERGEFORMAT 



Clock settings: Users can make change when need to calibration time.

[image: image13.png]R e e N Huazheng



 SHAPE  \* MERGEFORMAT 



Press “      ” key to exit this interface after the setting is done.

③Start Testing:

Press the “Start Test” key to enter the following interface:

[image: image8.png]State: Standing,countdown end in...

14 28

Minute Second

O O

Max.kV Test Count Volt.rate |

8 O kV. ]. O times 2 . 5kV/s

Stir Time Stand Time | Rest Time

10sec| 15w 5 wmin
‘ Nol.A”  No2.\/ No3.//




        

Fig.3 Standing

When the time for the oil sample to stand still is up, the instrument will automatically start to boost and detect the breakdown value of the oil sample.
[image: image14.png]R e e N Huazheng



[image: image15.png]Huazheng



we can click the "     " button to skip this process and directly enter the test process, in the countdown process of oil sample static placement.

Click the "     " key, and return to the main interface.

[image: image9.png]State: Voltage is Boosting...
Mean: 0. 00kV Cup 1
1] 35. 5kV

Max.kV Test Count Volt.rate |

8 O kV. ]. O times 2 . 5kV/s

Stir Time Stand Time | Rest Time

10sec| 15w 5 wmin
‘ Nol.A”  No2.\/ No3.//





Fig.4 Testing

[image: image16.png]


[image: image17.png]


[image: image18.png]


[image: image19.png]


During the step-up process, the instrument displays the voltage value in real time. Press the stop key to pause the voltage rise, press the "     " or "     " key to continue the voltage to rise, press the "     " key to reduce the voltage, and press the "     " key to end the test and return to the main interface

Test from the 1th Cup starts, and the interface shows "Voltage is boosting" and "Standing, countdown end in ...", according to priority finish 3 cups bit, again from 1th to begin testing, until have completed the user setting the boost times, buzzer beep, and enter the interface below.

[image: image10.png]State: Test Result Display
Mean: 12. 95kV Cup 1
1]13. 20kV 2]15:.50kV.

3]12.60kV  4]13.40kV
5]12.00kV ~ 6]11. 00kV

Std.Dev.:1. 39kV ~ CV.:0. 107

Max.kV Test Count Volt.rate |

8 O kV. ]. O times 2 . 5kV/s

Stir Time Stand Time | Rest Time

10sec| 15w 5 wmin
‘ Nol.A”  No2.\/ No3.//





Fig.5 Test Result Display

[image: image20.png]


[image: image21.png]J



[image: image22.png]


[image: image23.png]Press "     " and "     " buttons to display the test data of the tested oil sample in turn according to the cup position sequence；Press the "     " key to return the initial interface；Press the "     " key to return the initial interface.
④Data viewing and printing:

Press the “Data View” key to enter the following interface:

  [image: image11.png]Page:15/99
Mean:12. 95kV

1]113.20kV  2]15. 50kV
3]12.60kV ~ 4]13. 40kV
5]12.00kV ~ 6]11. 00kV

|Std.Dev.: 1. 39kV
D/Mean:(. 107

emperature:  26°C
umidity:  45%RH
€:2011-08-09 12:12:24





Fig.6 Data View
Select “     ” or “     ” and select the records to be printed and select “      ”.

V.Precautions

The selection of the oil sample before testing, placing and electrode distance shall meet relevant national and industrial standards.

The operators or other personnel are strictly forbidden to touch the casing after the power is switched on to avoid accidents.

The power shall be cut off immediately if any abnormal event is found during the operation.

New or newly cleaned oil cup should be breakdown 24 times before tested, oil cup with clean oil invasion of bubbles when without tested.

VI.Maintenance

This equipment shall not be exposed in moist environment.

Keep the oil cup and the electrodes clean. Fill the cup with fresh transformer oil for protection during its idle. Check the electrode distance and check the tightness between the electrode tip and electrode bar screw thread before the cup is used again.

Instrument oil Cup box high voltage electromagnetic switches are oil-filled insulation, there should be regular observation oil surface, through transparent switch shell, as oil surface at distances from the top of more than 10MM, should unscrew the plugs supplemented in line with the GB2536 25th of transformer oil.

VII.Oil Cup Cleaning Method And Common Fault Clearances

1. Oil Cup Cleaning Method

⑴ Wipe the electrode surfaces and bars again and again with clean silk cloth.

⑵ Adjust the electrode distance with standard gauge.

⑶ Use petroleum ether (other organic solvents are forbidden) to clean three times. Each time shall follow the bellow procedures:

①Pour the petroleum ether into the oil cup till the cup is 1/4～1/3 full.

②Cover the cup mouth with a piece of glass cleaned by petroleum ether. Shake the cup evenly for one minute with certain force. 

③Pour away the petroleum ether and dry the cup with a blower for 2~3 minutes.

⑷ Use the oil sample to be tested to clean the cup for 1~3 times.

①Pour the petroleum ether into the oil cup till the cup is 1/4～1/3 full. 

②Cover the cup mouth with a piece of glass cleaned by petroleum ether. Shake the cup evenly for one minute with certain force.

③Pour away the left oil sample and then the test starts.

2. Agitating Blade Cleaning Method.

⑴ Wipe the agitating blade again and again with clean silk cloth until fine particles are not found on their surfaces. It is forbidden to touch the surfaces with hands.

⑵ Use forceps to clamp the blade; put them into petroleum ether and wash.

⑶ Use forceps to clamp the blade and dry them with a blower.

⑷ Use forceps to clamp the blade; put them into the oil sample to be tested and wash.

3. Oil Cup Storage

Method 1 Fill the cup with good insulating oil after the test is finished and place it stable.

Method 2 Clean and dry the cup under the above procedures and then put it into a vacuum dryer.

Note: The oil cup and agitating blade shall be cleaned under the above procedures after the first test and tests with poor oil.

4. Common Fault Clearances

⑴Power light off, screen display off

①Check the plugging of power plug.

②Check the condition of the protector tube inside the power plug.

③Check the socket electricity.

⑵No punch through oil cup

①Check inserting of connectors on circuit board.

②Check contacting of cap high-voltage switch.

③Check attracting of high-voltage contacts.

④Check break of high-voltage line.

⑶Light display contrast

① Adjust the potentiometer on the circuit board.

⑷Printer failure

①Check plugging of printer power line.

②Check plugging of printer data line.

VIII.Packing List

	No.
	Item
	Qty

	1
	Main engine
	1

	2
	Oil Cup
	3

	3
	Test elbow
	6

	4
	Stirrer
	5

	5
	Standard gauge
	1

	6
	Tweezers
	1

	7
	Power Line
	1

	8
	Allen wrench
	1

	9
	Fuse
	2

	10
	Printing Paper
	2


Max.kV：10～80kV


Stand Time：0～15min


Rest Time：0～5min


Stirrer：OFF，10sec～120sec，Continuous


Test Count：1～10 times


Volt.Rate：0.5kV/s～5.0kV/s








� = 1 \* GB2 �⑴�：YES  NO  � = 2 \* GB2 �⑵�：YES  NO  � = 3 \* GB2 �⑶�：YES  NO





Automatic Printing：YES NO





2010-07-06 08:08:45









