[image: image11.png]Timer





[image: image8.png]R e e N Huazheng




HZNQ-1102 

Oil Pour Point Tester
[image: image1.png]



Huazheng Electric Manufacturing (Baoding) Co., Ltd
Contents

1

I.Overview

1

II.The Working Conditions of The Instrument

1

III.Main Technical Specifications and Parameters

1

IV.Instructions

3

V.Pour Point Operation Method

4

VI.Solidification Point Operation Method

5

VII.Cloud Point Operation Method

6

VIII.Matters Needing Attention

6

IX.Packing List



I.Overview
1.Oil condensation point test tester has the characteristics of exquisite structural design and unique manufacturing technology, and is the latest product updated by modern large-scale industrial production test and analysis means. It conforms to GB/T 3535 and GB/T510 standards and is used to determine the pour point of petroleum products.

2.This instrument adopts compressor refrigeration technology, saving energy, simple operation, long service life, and its main technical indexes, such as cooling rate and cooling capacity, have exceeded similar products abroad.

3.Environmental conditions for use of the instrument: ambient temperature 5℃～30℃, humidity ≤85%, operating power AC220V±10%, 50Hz.

II.The Working Conditions of The Instrument

1.The instrument should be placed on a flat and firm work table, and air convection should be minimized in the environment.

2. Flammable materials may be involved in the use of this product, and no open fire (or devices that generate sparks) shall be allowed around during operation.

3. The instrument should be placed in a dry and ventilated place when not in use.

4. The power supply is single-phase three-wire system, must have a good grounding end.

III.Main Technical Specifications and Parameters

1. Refrigerant: R13 R502

2. Cooling time: ≤100min (ambient temperature ≤30℃)

3. Downtime interval: > 30min

4. Stirring motor power: 22W X 2

5. Allowed temperature rise of stirring motor: 80℃

6. Power: 800W

7. Voltage: 220V±10%,50Hz±1Hz

8. Temperature control range: room temperature ~ -70℃

9. Temperature control accuracy: ±0.5℃

10. Temperature display: LED

11. Temperature sensor: Pt100

12. Timer: LED

IV.Instructions

1、Preparation before testing

1. Before doing the test, please carefully read the standards related to this test. The relevant standards are: People's Republic of China Standard GB/T3535-2006 "Petroleum Products Pour Point Determination Method", GB/T510, understand and familiar with the standard Test methods, test procedures and test requirements set forth.

2. Prepare various test instruments and materials for testing according to the requirements stipulated by the standards.

3. Check the working status of the instrument to meet the working environment and working conditions specified in the manual.

4. Check the power supply and have a good grounding terminal; check the outer casing of the instrument and it must be in a good grounding state.

5. Inject a proper amount of absolute ethanol (industrial alcohol) into the bath, and it is advisable to put the test tube into the bath bath without overflow.

2、Instructions

1. Turn on the power switch, the temperature control table displays the bath temperature, turn on the stirring switch, the bath is stirred to start working, turn on the cooling switch, and the refrigeration unit works.

2. Set the bath temperature point to the desired value according to the needs of the test project.

Temperature control table setting method: The upper row shows the PV red word as the temperature measured by the bath, and the lower row shows the SP green word as the temperature control point value to be controlled. Press the “[image: image2.emf]” function key once and press the ▲ or ▼ key to display the lower row as the desired temperature control point value. Press the “[image: image3.emf]” function key again to return to the standard mode. The temperature control meter is automatically constant at the set temperature point.

3、Timer Settings

[image: image9.png]R e e N Huazheng




1.The setting of the timer time is set by the "+" and "-" keys. The middle "S" indicates the second, the "M" indicates the minute, and the "H" indicates the hour (via the middle "+" and "-" keys. Select "S", "M", "H"). For example, "60S12" marks 60 seconds 12 seconds.

2.When the “Timer” switch is turned on, the timer starts counting. When the number on the LED matches the set number, the buzzer will start to beep, indicating that the time is up. At this time, the “Timer” switch is turned off to turn off the buzzer.

3.The pour point is tested in accordance with GB/T3535-2006 "Peak Point Determination Method for Petroleum Products". For detailed methods, see GB/T3535-2006 "Pour Point Determination Method for Petroleum Products".

4. If the actual value detected by the glass thermometer does not match the displayed value of the temperature control table, correction is required. For example, if the meter display value is -30.0 °C and the glass thermometer detection value is -29.7 °C, press the "[image: image4.emf]" button for more than 4 seconds, the PV window displays rE, and then press the "[image: image5.emf]" button. The PV window displays rt to make the SP value. It is 0.3 (if the glass thermometer has a detection value of -30.3 °C, the SP value is -0.3). After the correction is completed, press the "[image: image6.emf]" button for 4 seconds or more to exit (the key is automatically returned to the standard mode after 1 minute without pressing the button. (See the specification of the temperature control table for details).

Warning: when the instrument fails, the power should be cut off immediately. After the cold groove of the instrument is restored to room temperature, please ask professional personnel to repair and remove the fault before continuing to use, to prevent accidents!V.Common Troubleshooting

	Fault phenomenon
	Cause Analysis
	Method of exclusion

	Power switch does not work
	(1) Fuse blown

(2) The power cord is loose or the plug is not in good contact.
	(1) Replace the fuse

(2) Exclude after checking the power supply and plug

	Temperature display is confusing or not displayed
	(1) Sensor plug contact is not good

(2) Sensor damage

(3) Internal circuit failure
	(1) Re-inserted

(2) change sensor

(3) Ask the professional to repair

	Cooling rate is slow or not cooling
	(1) too much dust in the condenser

(2) Refrigerant leakage or circuit failure
	(1) Open the front door and side door of the refrigerator to remove dust

(2) Ask the professional to repair


V.Pour Point Operation Method

1. Name: Pour Point

2. Method summary: After the sample is cooled, check the fluidity of the sample every 3°C, and record the lowest temperature at which the sample can flow as the pour point.

3. Operation method

①Pour the sample into the dry and clean test tube to the ring mark, fix the thermometer in the center of the test tube with a stopper, make the bottom of the thermometer bubble 8 mm～10 mm from the bottom of the test tube, and then inject 1 mL～ into the bottom of the sleeve. 2 mL of absolute ethanol.

Then put the test tube into the sleeve to form the test tube assembly.

②Immerse this test tube assembly in the prepared cooling bath. From the first observation of the fluidity of the sample, the fluidity of the sample should be observed every time the temperature drops by 3°C, paying special attention not to disturb the lumps in the sample. The thermometer should not be moved after the sample is cooled enough to form paraffin crystals. When the temperature is low, the condensed water mist will hinder the observation. You can wipe the test tube with a cleaning cloth moistened with a wiper close to the temperature of the cold bath to remove the outer surface Water mist.

③When the test tube is tilted and the sample does not flow, immediately place the test tube in a horizontal position for 5 seconds (measured with a timer), and carefully observe the surface of the sample. If the sample shows any movement, immediately put the test tube back into the bath Or in the casing, when the temperature is lowered by 3°C, re-observe the fluidity of the sample.

④ Operate in this way until the test tube is placed in a horizontal position for 5 seconds and the sample in the test tube does not move. Record the thermometer readings observed at this time, and add 3°C to the recorded result as the pour point or lower Pour point.

VI.Solidification Point Operation Method
1. Name: solidification point

2. Summary of the method: Put the sample in a prescribed test tube and cool it to the expected temperature. Tilt the test tube to 45° from the horizontal and let it stand for 1 minute. Observe whether the liquid level moves. Use the highest temperature when the liquid level does not move. As the freezing point of the sample, it is expressed in °C.

3. Operation method

①Pour the sample into the dry and clean test tube to the ring mark, and fix the thermometer in the center of the test tube with a stopper, so that the bottom of the thermometer bubble is 8 mm to 10 mm from the bottom of the test tube. Then inject 1 mL to 2 mL of absolute ethanol into the bottom of the sleeve.

Then put the test tube into the sleeve to form the test tube assembly.

② Immerse the test tube assembly in the prepared cooling bath. The temperature of the cooling bath should be 7 ℃ ~ 8 ℃ lower than the expected freezing point of the sample.

③When the temperature of the sample cools to the expected freezing point, tilt the sleeve and test tube assembly immersed in the cooling bath to 45° from the horizontal and keep it for 1 minute. At this time, the sample is required to be still immersed in it. Inside the coolant.

④Then carefully remove the test tube assembly from the cooling bath, quickly wipe the outer wall of the sleeve with absolute ethanol, place it vertically, and observe through the sleeve whether the liquid level in the test tube shows signs of movement, until it is determined that a certain test temperature can make the sample When the liquid surface stops moving and rises by 2 ℃ and the sample liquid surface can move again, the temperature at which the sample liquid surface does not move is taken as the freezing point of the sample.

VII.Cloud Point Operation Method
Name:cloud point

Under specified conditions,maximum temperature at which clear and transparent liquid petroleum products,biodiesel and biodiesel blend fuels first appear foggy or turbid due to the emergence of wax crystals

2.Method summary:put clear and transparent sample into the instrument,Cool the sample by step.Judge whether the sample has wax crystal formation by visual observation or continuous monitoring of optical system. The highest sample temperature when wax appears at the bottom of the test tube for the first time
3.Operation method
Inject the sample into the dry and clean test tube to the annular marking, fix the thermometer in the center of the test tube with a plug, and the thermometer and the test tube are on the same axis, and the mercury ball of the thermometer just contacts the bottom of the test tube.

The gasket and casing shall be clean and dry. Put the gasket on the top of the casing. Before putting the test tube, the gasket and casing shall be put into the cooling medium for at least 10 minutes. When the empty casing is cooled, the cover can be covered. Then put the test tube into the casing. It is not allowed to put the test tube directly into the cooling medium

When the thermometer reading of the observation tube drops by 1 ℃, take out the tube quickly without stirring the sample, observe the cloud point, and then put it into the casing. The whole process shall not exceed 3S

When the sample is cooled to 9 ℃ and there is no cloud point, move the test tube into the second bath sleeve with the temperature maintained at - 18 ℃ ± 1.5 ℃. The sleeve cannot be transferred during the transfer of the test tube. If the sample is cooled to - 6 ℃ and there is no cloud point, move the test tube into the third bath sleeve with the temperature maintained at - 33 ℃ ± 1.5 ℃. In order to determine the very low cloud point, it is necessary to convert directly in two baths, and the bath temperature of each bath should be consistent with the figure below. In all cases, if the sample in the test tube has no cloud point and the bath temperature currently used by the identified sample reaches the lowest temperature in the corresponding sample temperature range in the figure below, the test tube shall be transferred to the next cold bath

[image: image7.png]Sample Temperature (ST)/C Casing Temperature/C
ST>+10 0+0.5

+10=>ST>—7 —17.2+0.5

—7=ST>—24 —34.240.5

—24>ST>—11 —51.2+0.5

—41>8T>~58 —68.240.5

—58=ST>~15 —85.2+0.5





5.When wax crystals appear at the bottom of the continuously cooled test tube, record the thermometer reading as cloud point
VIII.Matters Needing Attention

1.The instrument should be placed in a ventilated and dry place and kept clean.

2.The instrument power supply must be well grounded.

3.The tilt of the instrument should not be greater than 45° during handling and installation.

4.Do not dry heat heater without medium.

5.If you need to start again after turning off the refrigerator, you must do so at an interval of 5 minutes, otherwise it will damage the compressor.

6.In case of any malfunction of the instrument, professional personnel should be invited to repair or inform the company. Do not disassemble or disassemble the instrument randomly.
IX.Packing List

	No.
	Item
	Qty

	1
	Pour point in vitro
	2

	2
	Freezing point test tube
	4

	3
	Copper bush
	2

	4
	lid
	4

	5
	Pour point thermometer
	2

	6
	congelation point thermometer
	2

	7
	Black rubber plug
	4

	8
	White rubber plug
	2

	9
	Temperature sensor
	1

	10
	Fuse pipe
	4


1

[image: image10.png]Huazheng



[image: image11.png]