[image: image1.png]




HZNZ-100 

Battery Intemal Resistance Tester
   [image: image74.png]Wi

[EBE V) HERE @) ERFEE 00 [HE WA iR
2] 1ow BT
1 os6] 4108 w[FERL [EhEEUR





Huazheng Electric Manufacturing (Baoding) Co.,Ltd
Dear user:

Thank you for choosing HZNZ-100 Battery Intemal Resistance Tester
We hope that this instrument can make your work easier and more enjoyable, so that you can get the feeling of office automation in the test and analysis work.

  Before using the instrument, please read this manual, and operate and maintain the instrument according to the manual to prolong its service life.

"Just a light press, the test will be completed automatically" is the operating characteristics of this instrument.

     If you are satisfied with this instrument, please tell your colleagues; if you are not satisfied with this instrument, please call (0312) 6775656 to tell you to serve you at all times-Baoding Huazheng Electric Manufacturing Co., Ltd., our company will definitely make you satisfied !

Contents
1I.Summarize


3II.Junction Port And Cables Instruction


5III.Operation Rules


20IV.PC Software Instruction


23Chapter V Maintenance And Warning


25VI.Packing List




I.Summarize

1.1 General

This accumulator internal resistance tester utilized colour resistance touch screen. It can be operated via finger or the touch pen with convenience and flexibility. It is portable and the data can be known soon exactly. This instrument can store and process data of battery voltage and internal resistance. It will alarm for any battery failure. We can check data, download data and delete data on the instrument. Software in PC can display and .

1.2 Main function

Store 999 data. Users can check, download and delete data.
Storage way: Internal storage or SD card (It’s unnecessary to download data).
Download data. Transfer data to computer by SD card.
Humanized operation. Easy and clear, every step will display tips for reference.
Powerful data manage software. Friendly board. Provide the data management, printing, analysis, report, etc.
1.3 Main feature

High accuracy online test, large volume data
Instrument test range is 0.000mΩ-99.999mΩ.
Store 999 units battery information permanent （Each unit including batteries up to 500 pieces）.
Each report can indicate 500 pieces batteries information permanently.
Battery volume test range: 5AH-6000AH.
5”colour touch LCD. Chinese information modular operation. Great conversational interface.
Instrument can display chart and analyze column charts.
It can analyze volume and verify the battery quality level-better, good, bad.

It will display wave. Tell us practical battery highest, lowest and average power. Power wave (optional).
Via SD card port all test data can be stored in PC permanently. Track and analyze the battery “medical record”.
Powerful data management. The instrument can be used separately without computer.
Strong overload protection on power. Make it work more reliable.
Recover overcurrent fault protection create more convenience.
Utilized latest SOC chip simplified circuit increase the reliability.
Large volume battery and adapter supply power together to convene user.
Smart warning on low power ensure test accuracy .
Small foot print, light weight .
Automatic test mode convene user to measure.
Perfect PC test data analysis software.
Automatic analysis on battery “worse” status.
Record the history repot describe the battery quality curve.
Compare and analyze battery in same group.
Manage all batteries by different grades (Better, good, general, bad).
1.4 Specification

	Measure range
	Internal resistance: 0.000mΩ-99.999mΩ
Voltage: 0.000v--25v

	Min. resolution
	Internal resistance: 0.001mΩ
Voltage: 1mV

	Measure accuracy
	Internal resistance: ±0.5%rdg 　±6dgt
Voltage: ±0.2%rdg　 ±6dgt

	Power source
	12V, chargeable Li-battery

	Consumption
	8 hours standby mode

	Storage volume
	64Mbit Flash + 8G SD card

	LCD
	5”colour touch screen


II.Junction Port And Cables Instruction

2.1 Instrument port

[image: image2.jpg]NC
SD card port

NC
Charging port

Test line port

Power ON/OFF




Fig 2.1  Instrument port

Test line port: Open the cap and insert wire here

SD card port: Insert the SD card here

Charging port: Charge the instrument here

Keys:

Move the direction keys to select certain function

0-9 number keys

Enter： Confirming. Start the  test

Re Test: Test again. Do a test again

ESC： Back to last step

2.2 Wiring instruction

2.2.1 Wiring Notices: 

During test please don’t put test clip on bolts cap to avoid affecting test accuracy but can put it on pole bar or connection strip.
If we can’t clip the test clips please turn to test pin (option part).
Battery reference internal resistance please refer to information from manufacturer. If no supplied information please refer to instrument one. It will affect evaluation.
Confirming test mode. Before test line connecting battery please confirm test mode firstly to ensure whole process smooth.
2.2.2  battery connection diagram

[image: image3.png](\iiii




Fig 2.2.2 Poles connection

[image: image4.png]it




Fig 2.2.2-1 Connect connection strip

[image: image5.png]Wi ==




Fig 2.2.2 -2 Connection strip test

2.2.3 Test battery pole

When we test one group battery please test one battery all poles and connection strip (if there is) completely then begin next battery test. We can select any way to test different poles of one battery but please make sure all battery direction are same. After testing one battery pole, if we select “test connection strip” test the connection strip then begin next test on battery pole.

Test connection strip (before testing strip please test pole)

When we select “test connection strip” please add the line according to above diagram.
Before it is going to testing a battery bank, it has to finish all testing of poles and connectors of one battery, then start to testing next battery. First, test the poles, and then test the connectors. When

finished testing of all poles and connectors of one battery, test the next cell battery poles and connectors as the same direction.

III.Operation Rules

3.1 Fast test procedure

First step:：Press power button long time and release it until enter into interface[image: image6.png]


.

Second step：setting：main interface “[image: image7.png]Test Setup



” —“[image: image8.png]New Batt



”, enter into new battery parameters setting interface.

[image: image9.png]2015 10 07

Set Parameters

RoomNo - 0001 Tech-ID 0001 PileNo. 0001
CellNum 0500 B.Brand USER CellType 2v
Cnormal 100Ah Ri-ref 1.000me | Ulowlim. | 2.100V
OperMode | Manual | PoleNum 2

Test Options ‘
TestMode Pole only
TestPlan Volt & Resistance

Input Options
U-float 0.00V I-float 0.000A Temp. 25T

... StartTest .|} ... B¢

Prompt * The touch screen. please click.





Fig 3.1 Set test specification

Set machine room, technician, battery number

Set battery quantity (need the test battery quantity)

Battery manufacture selection:

Select “USER” then set single battery voltage, volume, reference internal resistance according to battery specification. 

If select other battery manufacturer then set these specification via menu.

Minimum voltage setting. During test if voltage is lower than this level device will send alarm but won’t affect test.
Test mode: By manual or automatic 

If manual test each change on test clip will request us to click “Start test” for next test. 

If automatic test every time change on test clip the instrument will test automatically don’t need manual click.
Pole quantity setting. According to poles quantity to set.
Test mode: 

Test the connection strip resistance. 

Select “test connection strip” If don’t test select “test on pole”.
Test item: 

Test voltage and internal resistance together.
Just test voltage no internal resistance.
Test all voltage firstly then test internal resistance.
It’s unnecessary to set float charging voltage, current, temperature won’t affect test.
Third step：Finish setting. Connect instrument and test line. Clip the test clips on battery (pay attention on pole. Red is positive and black is negative) Click “Start test” Click “Yes”.

Forth step: test process  

After finishing test of first battery, remove test clip to next battery( firstly remove test clip on the negative pole and then remove test clip on the positive pole ; Or remove two clips simultaneously ).
Oper Mode: 

1.  “ Manual”, click “Start “ test next battery .
2.  “ Auto” the instrument test automatically after 1-2s , user don’t click “ start “ .
[image: image61.png]R e e N Huazheng



Fig.3.1-1 Start test

Fifth step: end of test: Finish the test. Click “Return” to the main interface. If the whole battery group test was not finished when return but want to go on test click “Test battery” the test will be continued as per last test.

Sixth step: data check : “Data Mgmt” Select certain files and click “View” We’ll find detail data.
Seventh step: Generation of report: Print test report. External storage. SD card can transfer data to  computer by card reader. Just open the software in the computer will be ok. If it’s internal storage click “Data Mgmt” and click “Exp ALL” the data will be transferred to SD card. Then repeat first operation we’ll read all information.
3.2 Test parameters setting
[image: image10.png]Test Setup Continue Test
ab B
Data Mgmt System Mgmt




Fig.3.2 Instrument main interface

Specification setting: Set specifications of battery to be tested.
Test battery: Continue test of last un-finished battery group.
Data: management: Check and analyze test data.
System management: Including time setting, touch accuracy calibration, reference internal resistance maintenance, etc.

3.2.1 New battery specification setting

Main interface” [image: image11.png]Test Setup



” —“[image: image12.png]New Batt



” enter into parameters setting interface .
[image: image13.png][ @7 L= |

New Batt 01d Batt ESC




Fig.3.2.1Specifications setting

New battery specification: Custom setting on battery test specification.
Existing battery specification: Specifications of former test battery.
[image: image14.png]Set Parameters
RoomNo - 0001 Tech-ID 0001 PileNo. 0001
CellNum 0500 B.Brand USER CellType 2v
Cnormal 100Ah Ri-ref 1.000me | Ulowlim. | 2.100V
OperMode | Manual | PoleNum 2

Test Options
TestMode Pole only
TestPlan Volt & Resistance

Input Options
U-float 0.00V I-float 0.000A Temp. 25T
... Start Test ... EC ... ..

Prompt : The touch screen. please click.





Fig.3.2.1-1New battery specification setting

Machine room number: 0001-9999.
Technician number: 0001-9999.
Battery group number: 0001-9999.
Battery quantity: 1-500(The continuous test can be made for 500 pcs battery at most).
B.Brand:

Select USER, set the single battery voltage, volume, reference internal resistance specification according to actual demand.
When we select battery manufacturer, input voltage, battery volume setting via selection orderly the reference internal resistance will display automatically but also can be input by hand.
Single voltage: Rated voltage of battery.
Battery volume: Battery rated volume.
Reference internal resistance: Refer to battery manufacturer instruction.
Minimum voltage setting. During test if voltage is lower than this level device will send alarm but won’t affect test.
Test mode:：

By manual or automatic. If manual test each change on test clip will request us to click “Start test” for next test.

Automatic test. Every time change on test clip the instrument will test battery automatically don’t need manual click.
Pole quantity: 2/4/6/8

Test mode: 

Test on pole: Just test pole internal resistance.
Test the connection strip: Test pole internal resistance then test the resistance of pole connection strip.
Test item

Test voltage and internal resistance together.
Just test voltage no internal resistance.
Test all voltage firstly then test internal resistance.
It’s unnecessary to set float charging voltage, float charging current, battery temperature won’t affect test.
Start test: After setting specifications the clip instrument test clips on battery and click “Start test”.
Return: Return to last interface.
3.2.2 Existing battery specification

Main interface” [image: image15.png]Test Setup



” —“[image: image16.png]q§7

01d Batt



”, choose setting parameters 
[image: image17.png]ID. B.Brand RoomNo. CellType Cnormal Ri-ref
1 2V-1 0001 2V 100Ah 1.000m
2 2V-1 0001 2V 100Ah 1.000m
3 12v1 0001 12v 10Ah 18.700mQ
4 12v1 0001 12v 10Ah 18.700mQ |
5 12v1 0001 12v 10Ah 18.700mQ
6 2V-1 0001 2V 100Ah 1.000m
7 2V-1 0001 2V 100Ah 1.000m
8 2V-1 0001 2V 100Ah 1.000m
9 USER 0001 2V 100Ah 1.000m
lRelAll___Del UnPage Downfage ViewSel P3G |




Fig 3.2.2 Existing battery specification

“Delete all” : Click it will delete all modular record.
“Delete”: Select certain record and then click it will delete present record.
Last page, next page: Find out record.
Check setting: Select certain record click “ check setting” enter the specification setting board.

3.3 Start Test

Parameters setup interface click “[image: image18.png]Start Test



”, ready enter into test interface.

[image: image19.png]Set Parameters
RoomNo - 0001 Tech-ID 0001 PileNo. 0001
CellNum 0500 B.Brand CellType 2V
Cnormal In i m. | 2.100V
Opertlode External storage. please check

the test line. start?
TestMode
TestPlan oYes oo ce
Input Options

U-float 0.00V I-float 0.000A Temp. 25T
... StactTest | ... EC

Prompt :

The touch screen, please click.




Fig.3.3  Start test

Connecting test line.
Select “Yes” enter the test interface.
3.3.1 test interface

Test item is different from test mode , different interface, see below figure . 

[image: image20.png]24.4c¢  34% 10-13-2015 10:40 H

2.108v

TestState Finish
OperateMode Manual
No. 0001
TestPlan Volt-Test
‘ Voltage 2.108V ‘

G (O Gimse)) ()





Fig. 3.3.3 Test voltage and internal resistance together
[image: image21.png]24.6c¢  33% 10-13-2015 11:01

0.615m0

TestState Finish
OperateMode Auto
No. 0001
TestPlan Ri-Test
Voltage 2.108V
Ri 0.615mQ

G G Gos)) ()
Prompt:  The touch screen please click.





Fig.3.3.3-1 Just measure voltage

[image: image22.png]0. 614n0

TestState Finish
OperateMode Manual
No. 0001
TestPlan Ri-1th pair pole
Voltage 2.108V
Ri 0.614mo

G (O Gimse)) ()





Fig 3.3.3-2 More than 2 poles voltage and internal resistance test

[image: image23.png]B ) =
Prompt:  The touch screen. please click-




Fig 3.3.3-3 More than 2 poles and connection strip test on voltage and internal resistance

[image: image24.png]0.614n0

TestState Finish
OperateMode Manual
No. 0001
TestPlan Ri-1th pair pole
Voltage 2.108V
Ri 0.614mo
Re |  —

G (O Gimse)) ()





Fig. 3.3.3-4 All battery group after tested

Test status: In the test or soon to finish.
Test mode: Select manual test or automatic when set specification.
Battery number: Testing battery number.
Test item: The test result we will get.
Battery voltage: Present battery voltage.
Battery internal resistance: Present battery internal resistance.
Connection resistance: Test resistance between pole connection strip.
Re-test: Test battery again and replace last data when save new data.
If manual test then each change on test clip will request us to click “Start test” for next test. If automatic test every time change on test clip the instrument will test automatically don’t need manual click.
3.4 Continue test

3.4.1 Continue test mode: unfinished test battery bank and  historical unfinished battery bank .
3.4.2  Unfinished test battery bank 

main interface” [image: image25.png]q§7

Continue Test



” enter into continue test
[image: image62.png]R e e N Huazheng




Fig 3.4.2  Test battery

Continue test please connect battery test trip and click “Yes” to start test.
Main interface” [image: image26.png]


” — “ choose test record” — “ View” —“[image: image27.png]DataTest



” —“[image: image28.png]Con-Test



”.
[image: image29.png]Start test No.003Confirm right
Connection!

. Yes . . No

Prompt * The touch screen. please click.




   Figure 3.4.2-1  test battery interface
3.5 Data management  

3.5.1 Data save: internal storage and external storage (SD Card ).
Main interface “[image: image30.png]


” , enter into data management interface.
3.5.2 External storage 
[image: image31.png]2015 11 11

1D. File name

1 R00001-01-151006141240.NZY
2 R00001-01-151006143155.NZY
3 R00001-01-151006143437 .NZY
4 R00001-01-151006144632.NZY
5 R00001-01-151006144913 .NZY
6
7
8
9

R00001-01-151006145142 . NZY
R00001-01-151006145327 . NZY
R00001-01-151006145348.NZY
R00001-01-151006145651 . NZY

- |uDelsall . Del .. UpPage. DowPage...View..ESC..|
Prompt:  The touch screen. please click-





Fig.3.5.2 External storage interface

Name the file: R machine room number-battery group number-test date.
Every page displays 9-record. Read them by clicking “Last page” “Next page”.
Delete all： Click it to delete all record.
Delete: Select certain record and then click it will delete present record.
Check inforamtion: Select certain record click “ check setting” enter the specification setting board.
[image: image32.png]2015 11 11

1D. File name

R00001-01-151006141240.NZY
R00001-01-151006143155.NZY
R00001-01-151006143437 .NZY
R00001-01-151006144632 . NZY
RO0001-01-151006144913 .NZY

RSO R R

RO0001-01-151006145327 .NZY
RO0001-01-151006145348.NZY
R00001-01-151006145651.NZY
| Del-A1l  Del  UpPage DownPage  View .. ESC |

© o -





Fig. 3.5.2-1 Record selection

Select certain record and click “Detail” read every battery detail information.
3.5.3  Internal storage

[image: image33.png]1D. RoomNo. PileNo. Test Time

1 0001 01 2015-10-06 14:12
2 0001 01 2015-10-06 14:31
3 0001 01 2015-10-06 14:34
4 0001 01 2015-10-06 14:46
5 0001 01 2015-10-06 14:49
6 0001 01 2015-10-06 14:51
7 0001 01 2015-10-06 14:53
8 0001 01 2015-10-06 14:53
9

0001 01 2015-10-06 14:56
Del-All Del Exp-AlLuE UgP DownP Vi ESC





Fig.3.5.3Internal storage interface

Internal storage data will be transferred to computer via SD card for analysis.
Delete all： Click it to delete all record.
Delete: Select certain record and then click it will delete present record.
Click ：Download all” the data will be transferred to SD card.

Download: Select certain record click “Download” to transfer it into SD card.
Click “Last page” “Next page” to read.
Check: Select certain record click “Check” to find every battery detail information.
3.5.4 check test data 
[image: image34.png]NO. Voltage Ri Cap Decidg
1 2.108V 0.615mQ 100% Best
2 2.108V 0.605mQ 100% Best
3 2.108V 0.609mQ 100% Best
4 2.108V 0.620mQ 100% Best

. ViewSet  Chart  Datalest UpPage  DownPage  ESC





Fig.3.5.4 Read information

Battery number: Testing battery number.
Battery voltage: Present battery voltage.
Battery internal resistance: Present battery internal resistance.
Connection resistance: Connection resistance value (Only test battery on pole there is no connection resistance).
Volume: Left volume percentage.
Conclusion: Better, good, bad, worse, problem.
“------”means this item has not been tested.
Keys：

Setting check: Set the check specifications.
Column chart: All data of test battery displayed via column chart. Including voltage, internal resistance, volume.
Data test：View detailed test data.
Last page/Next page to check other battery data.
[image: image63.png]R e e N Huazheng




Fig.3.5.4-1. Battery data column chart

Column selection: Select different column to check different information. Such as voltage, volume, etc.

If check certain battery information click the column of that battery will display value accordingly.     
[image: image64.png]Huazheng




Fig.3.5.4-2 One battery test data

Certain battery detail data and specification.
Continue test: If present battery group test was not finished click it for continue test.
Re-test: Test present battery again.
Last line/Next line to check other battery information.
[image: image65.png]Set Parameters

RoomNo. 0001 Tech-ID

0001 PileNo. 0001

CellNum 0500 B.Brand
In

Cnormal

OperMode | M External storage. please check
the test line. start?

TestMode

TostPLan e wadon

USER CellType 2V
n prompt m. 2.100V

Input Options

U-float 0.00V I-float

0.000A Temp- 25t

... Start Test ...

... ESC_

Prompt *

The touch screen, please click.



Fig. 3.5.4-3 Continue test

Battery group test was not finished. Start any battery test connect test strip. Clik “Yes” start the test.
3.6 Re-test

Retest the battery

Main interface” [image: image35.png]


” — “ choose test record” — “ View” —“[image: image36.png]DataTest



” —“[image: image37.png]ReTest



”.
[image: image38.png]Concrete Data

NO. 1

Voltage 2.108V

Start test No.001Confirm right
Connection!

. Yes . . No

Prompt * The touch screen, p}ease click.




Figure 3.6  the test data of retested battery

3.7 Time setting
Main interface “[image: image39.png]L B

System Mgmt



”— “[image: image40.png]Date&Time



”, enter into date and time setting.
[image: image41.png]Date&Time

L’
TP Calculate

Default Setup

System Set

=7

ESC




Fig. 3.7 System setting/data setting interface

3.8 Touch screen calibration

Main interface “[image: image42.png]L B

System Mgmt



” — “[image: image43.png]4

TP Calculate



” enter into calibration interface .
[image: image44.png]Default Setup ESC




Fig. 3.8 Touch screen calibration

Click “Yes” Touch screen calibration utilized 5 points mode click the “+” orderly. When finish calibration it will return automatically. If any calibration mistake touch screen is not smart we can enter the interface again to calibrate it.

3.9 Back light memory setting when it’s off 

Main interface” [image: image45.png]L B

System Mgmt



”—“ [image: image46.png]Svstem Set



”，enter into parameters setting.
[image: image47.png]Auto off time 10 min

Backlight time

Alarm Close

Storagelocation SD Storage

s SAVE i | s BSCa





Fig. 3.9  System setting

Automatic off time: Maximum value is 120 minutes, 0 means no automatic off.
Back light off time: Maximum value is 120 minutes, 0 means no automatic off.
Alarm volume: Close-voltage lower than minimum level won’t alarm; Open-alarm user.
Storage：

Internal storage. Test data will be stored internal it can store maximum 99 group. Transfer it to PC.
External storage. The data is held in SD card. Maximum is 999 group.
Finish setting click “Save”. 

3.10 Save setting ( see 3.9)

main interface” [image: image48.png]L B

System Mgmt



 —“[image: image49.png]Svstem Set



”, enter into parameter setting interface.
3.11 Reference internal resistance maintenance

“System management” – “ Reference internal resistance”).
[image: image50.png]1 2v 100Ah 1.000n0
2 2v 150Ah 0.828u0
3 2v 170Ah 0.75%0
1 2v 2004k 0.699m0
5 v 2508h 0.680n0 |
6 2v 3004h 0.648u0
7 2v 3504h 0.600n0
8 2v 4004k 0.500n0
9 2v 4204h 0.47%0
| AddB.Brand Data | Save |  ESC





Fig. 3.11 Reference internal resistance edition

If we need to add or modify internal resistance value form manufacturer please enter this interface.
Delete battery manufacturer: Click it to delete battery data from present manufacturers.
Add battery manufacturer: Click it will create a new manufacturer.
[image: image51.png]Ref value Maintain

B-Brand 21 | Del B.Brand|,Add B.Brand

Prompt : The touch screen. please click.




Fig. 3.11-1 Add new manufacturer

Battery manufacturer name contains 4 letter or number at most.
[image: image52.png]B.Brand BATT | Dol BBrand | Add.B.Brand
ID. CellType CAP Ri
1 2V 1.000m





Fig. 3.11-2 Add manufacturer data

Voltage type: Select it according to battery type.
Battery volume: Input battery volume.
Reference internal resistance: Input the internal resistance from manufacturer.
Save above specification in test data specification select manufacturer will display battery information automatically.
3.12 Initialization setting 

main interface “[image: image53.png]L B

System Mgmt



—[image: image54.png]@

Default Setup



—[image: image55.png]q§7

Recovery



, enter into initialization setting interface.
[image: image56.png]g

Recover ESC




Fig. 3.12 Initialization

Initialization: Initialize all specifications as factory setting.
Product information: The instrument info.

IV.PC Software Instruction
4.1 System setting

This software requests following basic requirement.

The PENTIUM/100% compatible computer or note book with standard USB port 

A hard disc drive (Such as C) holding 80M volume at least.
Over 64M volume.
The screen resolution is larger than 1024*768. 

Computer system is higher thanWINDOWS2000/NT/ME/XP/2003.                 

4.2 Software installation

Open the setup.exe and user do installation according to tips. Operate the software open the SD card data.
4.3 Open test data 

Open SD card data

[image: image57.png]=) T
BE

| R00063-01-0A03121643. 171

FLESE & [R00001-01-0303121528. WX v]
TEEB @ [WEH ~





Fig. 4.2 Open SD card data

4.4 data analysis

[image: image58.png]TWHEE | witE | s

R
SEAM 100000
HMEE 24
BEER 1000 2
HZE 2

BENB 1
WRFRE 2015/10/9

PR ERE L
WABE  FREE
HFAE 0 000/
HFHR 0,004
KEAE 0 0007
MR 25T
FHEE 2 10000
FHEE 20100





Fig. 4.2-1

[image: image66.png]est

Data Mgmt System Mgmt



 
Figure 4.4-1  Graphic Correlation of cell

[image: image67.png]15 11:08 f

1/1 ColumnType|Ri MMaximum NO.: 4 Ri: 0.620mo
EMinimum NO.: 2 Ri: 0.605mQ
MCurrent NO.: 1 Ri: 0.615mo
0.744mo.

0. 149mo|
Hdiv |

0.000mo.

N2 3 4
.UpGroup . DownGroup . UpPage  DownPage . ESC



Fig.4.4-2  Maintenance interface

Reference internal resistance: It can be re-set and restart the computer to analyze the test data again.
Test report: Click it to create excel report.
4.5 Generation of report 

[image: image59.png]EtRRA T

Bt
1LAMRRSEELER:

IER: 280% BI5: 60%-80% SHt: dos-6o% FELH. <%
2. Bt AP A R AR




Fig. 4.5  Test report

[image: image68.png]Voltage 2.108V
Ri 0.615mQ





[image: image69.png]Start test No.003Confirm right
Connection!

. Yes . . No

Prompt * The touch screen, please click.



Fig. 4.5-1   Test report
[image: image70.png]¥ 6: \TfF\ROD0D1-01-151009101638. KZY,
WHERY | e

2500
2.000
1.500
1.000
0.500-
0.000-

Wtz

SERIEV]




Fig. 4.5-2  Battery status summarize
Fig. 4.5-3 Test result
Chapter V Maintenance And Warning

5.1 Failed to be turned on

Check the connection between instrument and Li-battery. The voltage may be too low to open the device. We may need to replace or charge battery in the instrument.
Instrument screen will display battery volume in the center. When there is[image: image60.png]


 we need to charge the instrument. 

Charging:

Insert the adapter to battery charging port. Connect adapter power to AC power socket.
When adapter light is green means finish charging.
Note: Longest charging time is 6-8 hour please don’t charge it over.
If instrument doesn’t reflect anything failed to turn off please open back cover get off the battery and insert it again then turn on the device.
If touch screen is not smart please calibrate it in “System management”.
When screen display “Insert storage card again” please get off the SD card and insert it again or replace SD card.
If we don’t operate it for a long while during test the instrument will close LCD or whole device automatically. Click the touch screen or press any key will wake it up. We can set the back light and off time ourselves.
5.2 Display “Inspect batter clip”

Firstly ensure battery clips and pole are connected well and testing battery voltage should be not less than 1.5V.
5.3 Test pin probe bent or can’t withdraw

Replace the probe:
Clip the probe on sleeve top by pincher.
Warning: Don’t clip the sleeve holding probe to avoid damaging pin.
 Pull out the pin probe straightly.
Clip new probe into sleeve.
Push the probe to a soft surface. Such as cardboard until the probe reach the sleeve top.
VI.Packing List

	No.
	Item
	Qty

	1
	Main engine
	1

	2
	Power Adapter
	1

	3
	Test line
	1

	4
	SD card
	1

	5
	Card reader
	1

	6
	Touch pen
	1


PAGE  

[image: image71.png][ 6:\ TAF\RO00D1-01-151009101638. KZT. (A=

BE® Wt

1 J2.130 7

8 R FEr
2

SEAE G



[image: image72.png]E it ER R A

HUSHR:
HUSER:
REA:



[image: image73.png]% it g PR RR

HNERS 00001
o1 00001
2v 24 FBETRR 2. 1007 BAPFE|  1.000n0
Elo.000v| AR | 0.0004 | LB 0. 000V iR | 25T
2016-10-9 10:16 TFH9eaFE 2. 700mQ FHeaE| 20107
Walhin

@i PR it BRI
®HE 1RERERLEP o
R 1RERERLEP o
&G Fer 1 .
%t &L 0




