[image: image12.png]R e e N Huazheng




  [image: image23.png]


                           

HZ1491
SF6 Gas Density Relay Calibrator

    [image: image1.png]


 
Huazheng Electric Manufacturing (Baoding) Co., Ltd

Dear user:

Thank you for choosingHZ1491 SF6 Gas Density Relay Calibrator.

     We hope that this instrument can make your work easier and more enjoyable, so that you can get the feeling of office automation in the test and analysis work.

  Before using the instrument, please read this manual, and operate and maintain the instrument according to the manual to prolong its service life.

"Just a light press, the test will be completed automatically" is the operating characteristics of this instrument.

     If you are satisfied with this instrument, please tell your colleagues; if you are not satisfied with this instrument, please call (0312) 6775656 to tell you to serve you at all times-Baoding Huazheng Electric Manufacturing Co., Ltd., our company will definitely make you satisfied !

Contents

1I.Overview


1II.Technical Indicators


2III.Main Features


3IV.Interface And Operation




I.Overview
SF6 gas density relay is an important component used to monitor the change of SF6 gas density in the SF6 switch body during operation. Due to the infrequent action of SF6 gas density relay in field operation, it often has inflexible action, poor contact and other phenomena after a period of time, and some also show that the temperature compensation performance of density relay becomes worse, Sudden change of ambient temperature often leads to misoperation of SF6 density relay. Therefore, DL / T596-1996 preventive test code for power equipment stipulates that each SF6 switch user shall regularly calibrate the SF6 gas density relay.JKMD20 SF6 density relay calibrator is developed according to DL / t259-2012: Calibration Specification for SF6 density relay issued by the national energy administration.

The instrument adopts 7.0 inch color LCD touch screen for human-computer interaction, which is more convenient to operate. Built in high-capacity rechargeable lithium battery pack provides ultra long standby time. The standard USB interface is configured, and the historical data can be exported to the USB flash disk through the interface.

In addition to the calibration of SF6 density relay, the instrument also has the function of calibrating pressure gauge and density gauge. The instrument can automatically detect abnormal conditions during measurement. In case of pipeline leakage, air leakage and insufficient inlet and outlet air pressure, the instrument will automatically stop operation and pop up a warning.

In a closed container, the SF6 gas pressure at a certain temperature can represent the density of SF6 gas. Traditionally, 20 ℃ is the SF6 gas pressure as the standard value. During on-site verification, the measured pressure value shall be converted into its corresponding pressure value at 20 ℃ under different ambient temperatures, so as to judge the performance of SF6 density relay.

II.Technical Indicators
Verification object: alarm 1, alarm 2, locking 1, locking 2, overpressure alarm contact signal

Calibration range: 0 ~ 1MPa (absolute pressure)

Instrument accuracy: grade 0.15

Pressure display resolution: 0.0001mpa

Inlet pressure: 0.9 ~ 1MPa (relative pressure)

External interface: USB (supports USB mobile storage devices no larger than 4GB)

Working environment: temperature - 10 ~ 50 ℃ (the best operating temperature is 5 ~ 50 ℃)

Relative humidity: 5 ~ 90% RH

Display resolution: 800 × 480 (pixel) color touch screen

Storage capacity: 50 groups

Measured contact voltage: DC24V

Temperature acquisition mode: manual / automatic

Air inlet connection: φ 4 butt type quick plug

Air inlet pipe: standard configuration φ 4 / PFA pipe

Measurement connection: φ 4 butt type quick plug

Measuring pipeline: standard configuration φ 4 / PFA pipe

Exhaust connection: φ 4 quick connector

Exhaust pipe: Standard φ 4 / PFA pipe

Mobile power supply: rechargeable lithium battery pack, which can work continuously for 8 hours

Charging voltage: 100 ~ 240VAC, 50 / 60Hz

Overall dimension of host: 325mm long × 275mm wide × Height 140mm

Overall dimension of accessory box: 360mm long × 250mm wide × Height 110mm

Instrument weight: 4.5kg

III.Main Features
The density relay calibrator adopts high-performance single chip microcomputer for detection and control, and the core components adopt imported components, with high precision, good repeatability and high reliability.

The initial state of 1 ~ 5 groups of contacts of density relay is in any form (normally open or normally closed), and the response value and recovery value of 1 ~ 5 groups of contacts can be measured at the same time.

Equipped with color touch LCD screen, it has simple operation and beautiful interface, and all parameters and states are clear at a glance.

Automatically complete the pressure measurement and 20 ℃ value conversion, so as to complete the dynamic automatic compensation between pressure and temperature. It also displays the pressure under the measured ambient temperature, the pressure under the ambient temperature at 20 ℃ and the ambient temperature. The problem of difficult field calibration of SF6 gas density relay is completely solved.

All testing processes are automatically completed by single chip microcomputer without manual intervention, which avoids cumbersome manual air path adjustment operation.

It can store 50 test results at the same time, and has the function of power failure data protection. It can query and print the previous test results at any time.

There is no need for constant temperature chamber in the calibration process, and SF6 gas density relay and pressure gauge can be calibrated within any effective temperature range.

It is equipped with many types of transition joints. The density relay of most types of switches can be verified on site without disassembly.

In addition to the calibration function of density relay, the instrument also has the calibration functions of 20 ℃ pressure gauge and normal temperature pressure gauge.

The pressure reducing valve of gas cylinder is professionally designed, which can withstand 15mp (150kg), and the low-pressure output is adjustable. It is suitable for calibration in the metering room with nitrogen as the test gas.

IV.Interface And Operation
Matters needing attention

If only the battery is used for power supply, make sure that the battery power is displayed more than one grid before use.

All data in the interface attached to the manual are only used for demonstration and do not represent real test data;

During the calibration and measurement of densimeter and pressure gauge, the final result must be the inflation and pressure rise to the set air pressure. When the air pressure exceeds the set air pressure, click the "quick release" button to reduce the air pressure and then boost the inflation and pressure to the set pressure. If the final result is the pressure drop to the set pressure, the data deviation may be greatly inaccurate.

If there is a pressure problem during the test, the following warning interface will pop up: 

[image: image2.png]Warn

U Inlet and outlet pressure alarm

Note: 1.Please check the inlet pressure
2.Check the connection and sealing

of internal pipeline

Enter




The user needs to check whether the inlet pressure is sufficient and greater than the set pressure, whether the inlet and outlet pipelines are connected normally, whether there is air leakage, etc.

Measurement interface

Connect the locking, alarm and overpressure output terminals of the density relay with the calibrator


Connect red 1, red 2, green 1, green 2 and yellow 0 on the signal line carried with the instrument (one is in the form of crocodile clip and the other is in the form of banana head) to a terminal corresponding to the output of the locking, alarm and overpressure contact of the density relay to be measured (if there is such output, it will be connected; if not, it will be idle and unused). At the same time, Connect the five "black" crocodile clip or banana head terminals to the other terminal corresponding to the output of the locking, alarm and overpressure contacts of the density relay to be measured.

Note: the instrument can measure the relays of 5 groups of contacts at most at the same time, and the signal lines of the calibrator and the meter to be measured can be connected only after confirming that the control power supply of the density relay calibrator has been disconnected.

[image: image3.png]-Nom. Press:
( 0.50 MPa

Relay Number:

JDQO001

Tester:

Tester01

Temp. Set

Manual:
®) Auto:

e

Type Set
Absolute Press () Absolue Mixed Press
Relative Press (_)Relative Mixed Press,

Pt (MPa)
0.0000

0.0000
0.0000
0.0000
0.0000

P20 (MPa)
0.0000

0.0000
0.0000
0.0000
0.0000

Pt (MPa)
0.0000

0.0000
0.0000
0.0000
0.0000

P20 (MPa)
0.0000

0.0000
0.0000
0.0000
0.0000

Atm.Press: 102.0 kPa
Test Press: 0.0000MPa
Test Press: 0.0000MPa(20C)

STAIH'

Print Back





Relay number：click this area to pop up the keyboard, and you can enter the relay number. Numbers, symbols and upper and lower case English characters can be entered. The total length is no more than 8 English characters. After entering, click 【Enter】on the keyboard interface to complete the input.

Tester：click this area to pop up the keyboard and enter the tester's number. Numbers, symbols and uppercase and lowercase English characters can be entered. The total length is no more than 6 English characters. After entering, click 【Enter】on the keyboard interface to complete the input.

Rated pressure：enter the rated pressure of the relay here.

Temperature setting：manual mode or automatic mode can be selected. Selecting manual mode requires manual input of temperature value. After selecting automatic mode, the instrument will measure temperature and display through its own temperature and humidity module.

Type setting：first determine the type of relay to be inspected, and then select the option consistent with the type of relay to be inspected.

Ambient air pressure：the instrument automatically measures and displays the current ambient atmospheric pressure.

Real time pressure：the real-time pressure output by the instrument is displayed here.

Battery power display：when the power display is one grid or lower, it needs to be charged in time.

【START】button：click this button, the instrument will automatically measure and display the results. After clicking this button, it will change to "stop" button. If you click 【STOP】 button, the instrument will stop measuring and return to 【START】 status.

【PRINT】 button：click this button, and the instrument will print the current or latest measurement results through the printer.

【BACK】button: click this button and the instrument will return to the main interface.

Densimeter calibration

[image: image4.png]@' Densimeter No.: MDBO0001|  Tester: Tester01

Type Set: () Absolute Press () Relative Press () Absolute Mixed Press (©) Relative Mixed Press
Temp(C): (O Manual: 0.0 (®Auto: 23.5 Actual Press: 0.0000 Mpa Micro-Ré \‘

Preset Pressure  Actual Value True Pressure

01 0.120 | 0.0000 I  0.0000 | | Micro-Drop)

02 0.120 | 0.0000 [l  0.0000 Stop )

03] 0.120 | 0.0000 [l  0.0000

04  0.120 | 0.0000 [»"  0.0000 | | Print |

05 0.120 || 0.0000 I}  0.0000 \
Back )

Atmospheric Pressure:102.0kpa





Density meter No.：click this area to pop up the keyboard, and you can enter the number of density meter. Numbers, symbols and upper and lower case English characters can be entered. The total length is 8 English characters. After entering, click【ENTER】on the keyboard interface to complete the input.

Tester：click this area to pop up the keyboard and enter the tester's number. Numbers, symbols and upper and lower case English characters can be entered. The total length is 6 English characters. After entering, click "enter" on the keyboard interface to complete the input.

Type setting：first determine the type of density table to be inspected, and then select the option consistent with the type of density table to be inspected.

Temperature setting：manual mode or automatic mode can be selected. Selecting manual mode requires manual input of temperature value. After selecting automatic mode, the instrument will measure temperature and display through its own temperature and humidity module.

Pressure：the real-time pressure output by the instrument is displayed here.

Battery power display：when the power display is one grid or lower, it needs to be charged in time.

Setting 1-5：here, you need to input the five pressure values set by the density meter to be tested, from low to high.

Output 1-5：after clicking the arrow button "" above, the instrument will display the output pressure value corresponding to the set pressure value in this area.

Gauge pressure 1-5:after the set pressure detection of this item is completed, the pressure value of the tested densimeter needs to be manually input here.

【Micro-Rise】:click this button to increase the output pressure of the instrument by 1kPa.

【Micro-Drop】:click this button to reduce the output pressure of the instrument by 1kPa.

【Stop】:click this button and the instrument will stop boosting.

【Print】:after the densimeter is tested and the gauge pressure is manually input, click the button, and the instrument will automatically print the parameters and pressure values displayed on the interface.

【Back】:click this button to return to the main interface.

Pressure gauge calibration

[image: image5.png]Pressure Gauge No.:

OAbsolute Press @ Relative Press Actual Press: 0.0000 Mpa

YLBOOO1 |Tester: TesterO1

01

HHEHEN

Preset Pressure
0.100MPa

0.200Mpa
0.300Mpa
0.400Mpa
0.500Mpa
0.600Mpa
0.700Mpa
0.800Mpa

Actual Value
0.0000Mpa

0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa

True Pressure
0.0000Mpa

0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa
0.0000Mpa

Ambient Press

i

102.0k
2.>] Pa
Micro-Rise
300
o
Stop
Print
-

Back





Pressure gauge No.: click this area to pop up the keyboard and enter the number of the pressure gauge. Numbers, symbols and upper and lower case English characters can be entered. The total length is 8 English characters. After entering, click "enter" on the keyboard interface to complete the input.

Tester: click this area to pop up the keyboard and enter the tester's number. Numbers, symbols and upper and lower case English characters can be entered. The total length is 6 English characters. After entering, click "enter" on the keyboard interface to complete the input.

Typesetting: first determine the type of pressure gauge to be inspected, and then select the option consistent with the type of pressure gauge to be inspected.

Pressure: the real-time pressure output by the instrument is displayed here.

Battery power display: when the power display is one grid or lower, it needs to be charged in time

Setting 1-8: Here you need to input 8 pressure values set by the pressure gauge to be tested, from low to high.

Output 1-8: the instrument will display the corresponding output pressure value in this area.

Gauge pressure 1-8: after the set pressure detection of this item is completed, the pressure value of the tested pressure gauge needs to be manually input here.

Key 1-8: after clicking a key "", the instrument will boost the pressure according to the set pressure value with the same key serial number.

"Micro-Rise": click this button to increase the output pressure of the instrument by 1kPa.

"Micro-Drop": click this button to reduce the output pressure of the instrument by 1kPa.

"Stop": click this button and the instrument will stop boosting.

"Print": when the pressure gauge is tested and the gauge pressure is manually input, click the button, and the instrument will automatically print the parameters and pressure values displayed on the interface.

"Back": click this button to return to the main interface.

System settings

Proportional valve calibration

Note: the proportional valve calibration procedure is only used when the output pressure of the instrument deviates. All instruments have completed the proportional valve calibration before leaving the factory and tested without error. The new instrument does not need calibration by default.

[image: image6.png]SetValue  AD Value
(Mpa) Tn Out

0.000] o | 211]
0.050]| 225 252 |
0. 100]| 446 | 295 |
0-150]| 668 | 335 |
0:200]| 890 | 384 |
0.250]| 1110| 425

SetValue  AD Value
(Mpa) Tn Out

0.300]| 1329| 471]
0.350]| 1550| 512]
0-400]| 1769| 551
0. 450]| 1990| 598
0.500]| 2212| 639

0.550]| 2433 | 681]

SetValue  AD Value
(Mpa) Tn Out

0.600]| 2657| 728
0.650]| 2877| 769]
0.700]| 3099| 812]
0.750] 338| 856

0800 3543 eoo|
-





This interface will calibrate the inlet and outlet of the proportional valve. Click the blank button on the right of the set pressure value to pop up a window as follows:

[image: image7.png]SetValue  AD Value SetValue  AD Value SetValue  AD Value
(Mpa) T Out (Mpa) I Out (Mpa) I Out

Pressure Set: 0.100 Mpa
Valve Input Value AD Value

448 296
Note:

Operate in strict accordance with the instructions.
Please do not raise the pressure above the set

pressure. After over pressure, exhaust and operate
again





Note: when adjusting the air pressure to the set air pressure, the final result must be the inflation and pressure rise to the set air pressure. You can first click the "fast charging" button for rapid inflation. When it is close to the set air pressure, click the "stop" button, and then click the "micro charging" button to slowly adjust to the set air pressure. If the air pressure exceeds the set air pressure, click the "exhaust" button to reduce the air pressure and boost the inflation pressure to the set pressure. If the final result is to reduce the pressure to the set pressure, the data may be greatly inaccurate.

Set pressure: the set pressure value is displayed at the top of the window to compare whether it needs to be calibrated. If not, you may click the wrong key and need to exit for re selection.

AD value display: the ad values of inlet and outlet are displayed under this label, that is, the values to be calibrated.

"Fast charge" button: click this button and the instrument will quickly rise to near the set pressure value.

"Stop" button: click this button to stop boosting

"Micro charging" button: click this button to increase the intake ad value by one, and the output pressure increases gradually. Through this button, it will be more stable and accurate to adjust to the set pressure.

"Exhaust" button: after clicking this button, the intake ad value will quickly decrease to 0 and the output pressure will also change to zero.

"Save" button: when the output pressure is adjusted to the set pressure, click this button, and the instrument will save the calibration data of the current set pressure and display it to the proportional valve calibration interface.

"Close" button: click the cross key at the upper right corner of the pop-up window to close the pop-up window.
Pipeline exhaust

[image: image8.png]~“Scitings
R

Calibration Pipe Exhaust

Actual Pressure: 0.0000 Mpa
Time Set: | 2017-03-03 15:15:15





Click this button, the instrument will empty the output air pressure, and the pressure box below the screen will display the output pressure in real time.

Time setting

[image: image9.png]~“Scitings
R

Calibration Pipe Exhaust

Actual Pressure: 0.0000 Mpa
Time Set: | 2017-03-03 15:15:15





Click the time box on the right of time setting, and the time setting dialog box will pop up as follows:

[image: image10.png]éﬁgmings

Clbetnpasty ¢ | 2 | 3
4 5 6
7 8 9

Actual Pressure: 0.0000 Mpa

Time Set: | 2017-03-03 1¢





Just input the time to be modified and click "OK" to complete the time reset.

Data browsing

The instrument will automatically record the test data of 50 density relays. If it exceeds the limit, it will automatically delete the earliest record and scroll for storage.

Click the "data browsing" button in the main interface to enter the data browsing interface, and the following operations can be carried out:

[image: image11.png]’Data View

Relay No.: JDQ0001

Tester: Tester01

Testing Environment Test Up Down
Related Press| 0.50Mpa Point|  p¢ P20 Pt P20
Temperature| 23 51 [Lockil| 0.4021 | 0.4023 | 0.3915 | 0.3824
Atm. Press [ 102.1kPa - 0.4021 | 0.4023 | 0.3915 | 0.3824

Type Relative Press - 0.4021 | 0.4023 | 0.3915 | 0.3824
Time 12:41:58 - 0.4021 | 0.4023 | 0.3915 | 0.3824
Date 2015-08-14 - 0.4021 | 0.4023 | 0.3915 | 0.3824
+ Y BN





[image: image13.png]R e e N Huazheng



“      ”：Click this button and the instrument will display the previous data.

[image: image14.png]R e e N Huazheng



“       ”：Click this button and the instrument will display the next data.
[image: image15.png]R e e N Huazheng



“       ”：After inserting the USB flash disk, click this button, and the instrument will transmit the currently displayed data to the USB flash disk.
[image: image16.png]R e e N Huazheng



“       ”：Click this button and the instrument will print the current piece of data through the printer.
[image: image17.png]Huazheng



“       ”：Click this button and the instrument will delete the current data.

[image: image18.png]


“       ”：Click this button and the instrument will return to the main interface.
1

[image: image19.png]


[image: image20.png]


[image: image21.png]


[image: image22.png]


[image: image23.png]